## General Information
|  |  |  | Document #: | 10037183 | Revision: | W
1 | Purpose and Scope
The purpose of this analysis is to identify hazards associated with the intended use, as well as, reasonably foreseeable misuse of the medical device and any potentially resulting harms in accordance with EN ISO 14971.
This document covers the patient, all intended users, and the device itself during manufacture, use, servicing, transportation, and storage of the medical device.
2 | Assumptions:
This hazard analysis assumes that the PB560, PB540, PB520 and PowerPack shall only be operated by qualified, properly trained personnel (i.e. used by prescription only). The users of thePB560, PB540, PB520 and PowerPack include: clinicians, physicians, nurses, homecare providers, patients and patient’s families.
Per the Risk Management Plan, 10035479, risk management activities are retrospective to comply with EN ISO 14971:2012, and Remediation Plan for Risk Management Documents 10130032, and shall be carried out in accordance with MITG Risk Management Method GS10004411. Because this is retrospective, the columns in the Product Risk Analysis table below that call for the 'Pre-mitigation risk assessment' will not be populated. The 'Post mitigation' (risidual) risk assessment columns will be the only ones used because the mitigations have already been applied. The PB500 Risk Assessment Team considered the probability of both P1 and P2 occurring together without defining each other separately. The residual probability of harm was arrived by combining engineering design expertise with data from a 3 years period of PB500 customer complaints.
Within this file cells that have been shaded brown represent the start of the hazard. All causes, harms and mitigations listed below this line are appropriate for the hazard. P1, P2 and risk zones are correct for the row associated with the 'harm x severity' found on the same row. Grey shading has been used to block out unused cells and are considered equivalent to 'not applicable' for documentation purposes.
3 | Intended Users and Use Environment
| · Risk Management activities are retrospective to comply with EN ISO 14971:2012 [R5] and Remediation Plan for Risk Management Documents 10130032 [R10], and shall be carried out in accordance with MITG Risk Management Method GS10004411 [R1]. | Intended Users and Use Environment:
The ventilator is a restricted medical device intended for use by qualified, trained personnel under the direction of a doctor and may be operated by clinicians, physicians, nurses, homecare providers, patients and patient’s families.
| 3.2 | The intended use environments for the above personnel include:
The PB500 ventilator is intended for use in institutional, home, and portable settings. The ventilator is not for use as an emergency transport ventilator.
| 3.3 | Intended Patient Population:
The PB500 series ventilators are indicated for the continuous or intermittent mechanical ventilatory support of patients weighing at least 11 lbs (5 kg) who require mechanical ventilation delivered invasively or non-invasively.
4 | System Definition/Exclusions:
N/A
5 | Characteristics Related to Safety
| 5.1 | Primary Operating Function(s):
The PB500 ventilators are applicable for adult and pediatric patients who require the following general types of invasive (via endotracheal and tracheostomy tubes) or non-invasive (via mouthpiece, nasal mask or full face mask) ventilatory support, as prescribed by an attending doctor: • Positive Pressure ventilation • Assist/Control, SIMV, or CPAP modes of ventilation • Breath types including Volume Control, Pressure Control, and Pressure Support PB560 also supports volume control ventilation
| 5.2 | Frequently Used Function(s):
N/A
| 5.3 | Essential Performance:
Essential Performance for the PB560 ventilator is defined within RE00147193.
6 | Standard Revisions:
The standards referenced in the Risk Anaysis Chart correspond with the following standard revisions:
| EN 60601-1:2006 + A12:2014 |  | Medical Electrical Equipment - Part 1: General Requirements for Safety - Collateral Standard - Safety Requirements for Medical Electrical Systems
| EN ISO 10993-1:2009 |  | Biological Evaluation of medical devices Part 1: Evaluation and testing within a risk management process
| EN 10651-2:2009 |  | Lung Ventilators for Medical Use- Particular Requirements for Basic Safety and Essential Performance Part 2: Home Care Ventilators for Ventilator-Dependent Patient
| IEC 60601-1-11:2015 IEC 60601-1-11:2015 |  | Medical electrical equipment – Part 1-11: General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment
| EN 60601-1-6:2010 + A1:2014 IEC 60601-1-6: 2010 +A1:2013 |  | Medical Electrical Equipment - Part 1: General Requirements for Safety -2- Collateral standard: Usability
| EN 60601-1-8:2007 + A11:2017 IEC 60601-1-8:2006 + A1:2012 |  | General Requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems
| EN ISO 80601-2-55:2011 |  | Medical electrical equipment Part 2-55: Particular requirements for basic safety and essential performance of respiratory gas monitor
| EN ISO 80601-2-72:2015 |  | Medical Electrical Equipment - Part 2-72:Particular requirements for basic safety and essential performance of home healthcare environments ventilator-dependent patients
| EN 60601-1-2:2015 IEC 60601-1-2:2014 |  | Medical Electrical Equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic Disturbances - requirements and tests
7 | PB560 Worst case Testing
Rationale for Worst Case Tubing: The characteristics of the patient circuit impact ventilator’s performance in multiple ways by dynamic interaction with the ventilator’s pneumatic system as well as through thermodynamic and fluid mechanic phenomena pursuant to varying patterns and dynamic characteristics of gas flow through the patient-ventilator system. The accuracy of the measurement of timing and magnitude of pressure and flow changes for detection are dependent on tubing characteristics. The accuracy for delivered and exhaled tidal volumes are directly related to the respective accuracy of the tubing compliance volume during inhalation and exhalation. The ratio of tubing compliance to the patient’s respiratory compliance as well as the rate of circuit pressurization (function of flow rate and circuit resistance) are among the main factors influencing the accuracy of tubing compliance volume. Therefore, the worst-case patient circuit for volume-based ventilation is considered to be the one with highest resistance and the highest compliance. In general, higher circuit resistance poses more challenges for pressure-based ventilation. For testing, the worst-case patient circuit will be selected from amongst the specified breathing circuits based on their compliance and resistance properties. Therefore, the worst-case patient circuit for verification is considered to be the one with highest resistance and the highest compliance i.e. P/N: 5094000 and 5093900.
Rationale for Bacteria Filters (As per RE00164878): 351U5856 Bacteria Filters are considered least favorable due to highest resistance and volume combination of all standard bacteria filters recommended for use with the PB560.

## Annex C Questions
|  |  | Document #: | 10037183 | Revision: | W
Reference ISO 14971 | Question | Factors to Consider | Response (N/A = Not Relevant) | Potential Hazards(s) | Potential Hazardous situation(s) | Notes / Comments
C.2.1 | What is the intended use and how is the medical device to be used? | What is the medical device's role relative to: • diagnosis, prevention, monitoring, treatment, or alleviation of disease, | The Puritan Bennett™ 560 Ventilator is indicated for the continuous or intermittent mechanical ventilatory support of patients weighing at least 11 lb (5 kg) who require mechanical ventilation. The ventilator is a restricted medical device intended for use by qualified, trained personnel under the direction of a doctor. Specifically, the ventilator is applicable for adult and pediatric patients who require the following general types of invasive or non-invasive ventilatory support, as prescribed by an attending doctor: • Positive Pressure ventilation • Assist/Control, SIMV, or CPAP modes of ventilation • Breath types including Volume Control, Pressure Control, and Pressure Support The ventilator may be operated by: • Respiratory therapists • Doctors • Nurses • Homecare providers • Patient and patient’s families | Refer to RE00076775 Ventilation Hazards | N/A | N/A
• compensation for injury or handicap or
• replacement or modification of anatomy, or control of conception?
• what are the indications for use (e.g., patient population)?
Does the medical device sustain or support life?
Is special intervention necessary in the case of failure of the medical device?
C.2.2 | Is the medical device intended to be implanted? | Consider: the location of implantation, the characteristics of the patient population, age, weight, physical activity, the effect of aging on implant performance, the expected lifetime of the implant, the reversibility of the implantation. | N/A | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.3 | Is the medical device intended to be in contact with the patient or other persons? | Nature of the intended contact, i.e. surface contact, invasive contact, or implantation and, for each, the period and frequency of contact. • surface contact points; • invasive contact / contact through body orifices; • implantable contact (permanent / absorbable) | N/A None of the parts of the ventilator is in direct patient contact. A breathing circuit is connected to the ventilator. The breathing circuit does not come into contact with the patient directly. The ventilator may be used with an endotracheal tube, tracheal tube, face mask, nasal mask or prongs, or mouthpiece. |  | N/A | N/A
C.2.4 | What materials or components are utilized in the medical device or are used with, or are in contact with, the medical device? | Consider: • compatibility with relevant substances; | Clinicians may use a variety of interfaces to connect patients to the ventilator: nasal masks or full face masks; endotracheal or tracheotomy tubes. |  | N/A | N/A
• compatibility with tissues or body fluids;
• whether characteristics relevant to safety are known;
• is the device manufactured utilizing materials of animal origin? See Annex I and ISO 22442 Annex I – Biological Hazards ISO 22442-1:2007 = provides requirements and guidance on risk management related to the hazards typical of medical devices manufactured utilizing animal tissues or derivatives
C.2.5 | Is energy delivered to or extracted from the patient? | Consider: • the type of energy transferred; | N/A | N/A | N/A | N/A
• its control, quality, quantity, intensity, and duration;
• whether energy levels are higher than those currently used for similar devices.
C.2.6 | Are substances delivered to or extracted from the patient? | Consider: • whether the substance is delivered or extracted; | The ventilator delivers regulated air and/or oxygen to the patient via a medical pipeline system or from gas cylinders. It is supplied to the patient circuit through check valves and exits to the patient circuit close to patient pressure. This ventilator is not for use with anesthetic gases. | Refer to RE00076775 Ventilation Hazards | N/A | N/A
• whether it is a single substance or range of substances;
• the maximum and minimum transfer rates and control thereof.
C.2.7 | Are biological materials processed by the medical device for subsequent re-use, transfusion, or transplantation? | Consider: the type of process and substance(s) processed (e.g., auto-transfusion, dialysis, blood component, or cell therapy processing). | N/A | N/A | N/A | N/A
C.2.8 | Is the medical device supplied sterile or intended to be sterilized by the user, or are other microbiological controls applicable? | Consider: • whether the medical device is intended for single use or to be reusable; | N/A The ventilator is not a sterile device | N/A | N/A | N/A
• any packaging;
• shelf-life issues;
• limitation on the number of re-use cycles;
• method of product sterilization;
• the impact of other sterilization methods not intended by the manufacturer.
• Is the medical device properly marked to inform the user whether it is for single use or to be re-used? • Does the packaging clearly indicate any limitation of handling or shelf-life? • Does the accompanying document clearly indicate proper methods of and agents to be used for cleaning or sterilization, and inform about frequencies of cleaning?
• Simplicity of disassembling / reassembling and any USE errors connected to these?
C.2.9 | Is the medical device intended to be routinely cleaned and disinfected by the user? | Consider: • the types of cleaning or disinfecting agents to be used and any limitations on the number of cleaning cycles. | All external panels and surfaces should be cleaned before and after each patient use and as often as necessary to keep the ventilator clean. Additionally the ventilator should be cleaned the ventilator periodically, whenever it is soiled or dirty, before any maintenance operation, and before storing the ventilator. Approved Cleaning Solutions for Exterior Ventilator Surfaces: Mild dishwashing detergent 70% isopropyl alcohol (rubbing alcohol) 10% chlorine bleach (90% tap water) Glutaraldehyde Hospital disinfectant cleaners Hydrogen peroxide 15% ammonia (85% tap water) Ammonia-based household cleaners Household cleaners Instruction for the cleaning of the accessories is to follow the manufacturer’s instructions for cleaning the ventilator’s accessories and components, including the patient circuit. The exhalation block is intended for single use by a single patient . It may periodically be cleaned, but it cannot be disinfected or sterilized. To maintain good measurement quality when used continuously, clean the exhalation block periodically. The exhalation block should be changed every 4 months and cannot be reused with any other patient. | Refer to RE00076775 Ventilation Hazards | N/A | N/A
• The design of the medical device can influence the effectiveness of routine cleaning and disinfection. In addition, consideration should be given to the effect of cleaning and disinfecting agents on the safety or performance of the device.
C.2.10 | Is the medical device intended to modify the patient environment? | Consider: • temperature; • humidity; • atmospheric gas composition; • pressure; • light. | The ventilator modifies the patient's breathing environment. User-selectable ventilation modes are: • Assisted Controlled Volume (V A/C) • Assisted Controlled Pressure (P A/C) • Volume Synchronized Intermittent Mandatory Ventilation (V SIMV) • Pressure Synchronized Intermittent Mandatory Ventilation (P SIMV) • Continuous Positive Airway Pressure (CPAP) • Pressure Support Ventilation with apnea ventilation (PSV/ST) | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.11 | Are measurements taken? | Consider: the variables measured and the accuracy and the precision of the measurement results. Also, users have to be aware of frequencies at which measurements have to be initiated, parameters influencing the result, consumables needed, how to handle or interpret the results. Standard issues are legibility and exactness of displays. Neglecting routine maintenance also might cause wrong results. | The ventilator measures and monitors patient respiratory parameters. A trained user will need to understand, evaluate and adjust parameters.
C.2.12 | Is the medical device interpretative? | Consider: whether conclusions are presented by the medical device from input or acquired data, the algorithms used, and confidence limits. Special attention should be given to unintended applications of the data or algorithm.
C.2.13 | Is the medical device intended for use in conjunction with other medical devices, medicines, or other medical technologies? | Consider: identifying any other medical devices, medicines, or other medical technologies that can be involved and the potential problems associated with such interactions, as well as patient compliance with the therapy.
C.2.14 | Are there unwanted outputs of energy or substances? | • Energy-related Consider: noise and vibration, heat, radiation (including ionizing, non-ionizing, and ultraviolet /visible/infrared radiation), contact temperatures, leakage currents, and electric or magnetic fields. | There are no unwanted outputs of energy or substances under normal operating conditions. | Refer to RE00076775 Ventilation Hazards | N/A | N/A
• The adverse effects of noise, vibration, heat, and waste products (exhaust gases) on physiology and psychology of users and thirds.
• Substance-related Consider: substances used in manufacturing, cleaning or testing having unwanted physiological effects if they remain in the product.
• Other substance-related Consider: discharge of chemicals, waste products, and body fluids.
C.2.15 | Is the medical device susceptible to environmental influences? | Consider: the operational, transport and storage environments. These include light, temperature, humidity, vibrations, spillage, susceptibility to variations in power and cooling supplies, and electromagnetic interference. | Yes – temperature, pressure, humidity, vibration, electrical/magnetic field. | Refer to RE00076775 Ventilation Hazards
C.2.16 | Does the medical device influence the environment? | Consider:: • the effects on power and cooling supplies; | Yes – alarm noises/notifications. | Refer to RE00076775 Ventilation Hazards | N/A | N/A
• emission of toxic materials;
• the generation of electromagnetic disturbance
C.2.17 | Are there essential consumables or accessories associated with the medical device? | Consider: specifications for such consumables or accessories and any restrictions placed upon users in their selection of these. | Consumables and Replacement Intervals • Air Inlet Filter - (Foam + Fine Particle) Once a month or more often, depending on the extent of soiling • Inspirited Bacteria Filter See manufacturer’s recommendation • Patient Circuit - See manufacturer’s recommendation - Single use Single patient • FIO2 Sensor - 14 to 18 months or more often in case of persistent calibration failure • Exhalation Block 4 months(*) (and for each new patient) humidifiers, water traps • Oxygen solenoid valve - Every 15,000 hours of use •Turbine - Every 15,000 hours of use •Exhalation solenoid valve - Every 15,000 hours of use •Cooling fan - Every 15,000 hours of use •Clean and disinfect Insp block - Every two years •Measurements check and - Every two years •calibration - Every two years •Battery, lithium-ion 4.8 Ah memory - Every two years •Battery, lithium, 3V - Every two years •Buzzer PCBA - Every two years | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.18 | Is maintenance or calibration necessary? | Consider: • whether maintenance or calibration are to be carried out by the operator or user or by a specialist; | Each time the exhalation block or circuit is removed and reinstalled or after installing a new exhalation block, the exhalation flow sensor must be recalibrated before the using the ventilator. This process is automatic and does not require the use of a measurement device. Each time the FIO2 sensor is removed and reinstalled, and on a weekly basis, the FIO2 sensor must be recalibrated before using the ventilator. This process does not require the use of a measurement device. | Refer to RE00076775 Ventilation Hazards | N/A | N/A
• are special substances or equipment necessary for proper maintenance or calibration?
C.2.19 | Does the medical device contain software? | Consider: whether software is intended to be installed, verified, modified, or exchanged by the operator or user or by a specialist. | Software is installed by a qualified service engineer | N/A | N/A | N/A
C.2.20 | Does the medical device have a restricted shelf-life? | Consider: labeling or indicators and the disposal of such medical devices when the expiration date is reached. | N/A | N/A | N/A | N/A
C.2.21 | Are there any delayed or long-term use effects? | Consider: ergonomic and cumulative effects. Examples could include pumps for saline that corrode over time, mechanical fatigue, loosening of straps and attachments, vibration effects, labels that wear or fall off, long term material degradation. | The patient can become dependent on the ventilator to provide respiration. The ventilator will remain in good service condition when the preventative maintenance schedule is adhered to | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.22 | To what mechanical forces will the medical device be subjected? | Consider: whether the forces to which the medical device will be subjected are under the control of the user or controlled by interaction with other persons. | Shipping – ISTA 2A; shock and vibration The ventilator is suitable for use in institutional, home, and portable settings. It is not intended for use as an emergency transport ventilator. - suitable for use on commercial aircraft, per FAA requirements | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.23 | What determines the lifetime of the medical device? | Consider: aging and battery depletion. | The lifetime of the device is determined by a number of factors: • Life of product line is from inception to withdraw from market • Shelf-Life of components is determined by Stability and Packaging evaluation. | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.24 | Is the medical device intended for single use? | Consider: does the medical device self-destruct after use? Is it obvious that the device has been used? | No | N/A | N/A | N/A
C.2.25 | Is safe decommissioning or disposal of the medical device necessary? | Consider: the waste products that are generated during the disposal of the medical device itself. For example, does it contain toxic or hazardous material, or is the material recyclable? | The ventilator is marked to ensure that it enters the recycled waste system. 'Recycle electrical and electronic items with this marking according to local regulations (EN50419). Do not discard with ordinary household waste. This includes the ventilator and the battery.' | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.26 | Does installation or use of the medical device require special training or special skills? | Consider: the novelty of the medical device and the likely skill and training of the person installing the device. | The ventilator is for use as prescribed by a doctor. The ventilator may be operated by: • Respiratory therapists • Doctors • Nurses • Homecare providers • Patient and patient’s families It is the responsibility of the clinician or clinical educator to ensure that both the patient and the caregiver fully understand the daily use of the ventilator. | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.27 | How will information for safe use be provided? | Consider: • whether information will be provided directly to the end user by the manufacturer or will it involve the participation of third parties such as installers, care providers, health care professionals, or pharmacists, and whether this will have implications for training; | Training and manuals | Refer to RE00076775 Ventilation Hazards | N/A | N/A
|  | • commissioning and handing over to the end user and whether it is likely/possible that installation can be carried out by people without the necessary skills; |  | Refer to RE00076775 Ventilation Hazards | N/A | N/A
|  | • based on the expected life of the device, whether re-training or re-certification of operators or service personnel would be required. |  | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.28 | Will new manufacturing processes need to be established or introduced? | Consider: new technology or a new scale of production. | N/A | N/A | N/A | N/A
C.2.29 | Is successful application of the medical device critically dependent on human factors such as the user interface? | Factors that should be considered include user interface design features that can contribute to use error. Features should be designed so that they cannot be easily misused by busy users in an environment where distractions are commonplace, e.g. device control, symbols used, ergonomic features physical design and layout, hierarchy of operation, menus for software driven devices, visibility of warnings, and audibility of alarms, standardized color coding.
C.2.29.1 | Can the user interface design features contribute to use error? | Factors that should be considered are user interface design features that can contribute to use error. Examples of interface design features include: control and indicators, symbols used, ergonomic features, physical design and layout, hierarchy of operation, menus for software driven devices, visibility of warnings, audibility of alarms, and standardization of color coding. See IEC 60601-1-6 for additional guidance on usability and IEC 60601-1-8 for guidance on alarms. | Yes | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.29.2 | Is the medical device used in an environment where distractions can cause use error? | Consider: • the consequence of use error; | Yes | Refer to RE00076775 Ventilation Hazards | N/A | N/A
|  | • whether the distractions are commonplace; |  | Refer to RE00076775 Ventilation Hazards | N/A | N/A
|  | • whether the user can be disturbed by an infrequent distraction |  | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.29.3 | Does the medical device have connecting parts or accessories? | Consider: the possibility of wrong connections, similarity to other products’ connections, connection force, feedback on connection integrity, and over- and under-tightening | Yes | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.29.4 | Does the medical device have a control interface? | Consider: spacing, coding, grouping, mapping, modes of feedback, blunders, slips, control differentiation, visibility, direction of activation or change, whether the controls are continuous or discrete, and the reversibility of settings or actions. | Yes | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.29.5 | Does the medical device display information? | Consider: visibility in various environments, orientation, the visual capabilities of the user, populations and perspectives, clarity of the presented information, units, color coding, and the accessibility of critical information. | Yes | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.29.6 | Is the medical device controlled by a menu? | Consider: complexity and number of layers, awareness of state, location of settings, navigation method, number of steps per action, sequence clarity and memorization problems, and importance of control function relative to its accessibility and the impact of deviating from specified operating procedures. | Yes | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.29.7 | Will the medical device be used by persons with special needs? | Consider: the user, their mental and physical abilities, skill, and training, ergonomic aspects, the use environment, installation requirements, and the patient's capability to control or influence the use of the medical device. Special attention should be paid to users with special needs, such as handicapped persons, the elderly, and children. Their special needs might include assistance by another person to enable the use of a medical device. Is the medical device intended to be used by individuals with various skill levels and cultural backgrounds? | Only if the user is differently abled | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.29.8 | Can the user interface be used to initiate user actions? | Consider: the possibility of initiating a deliberate action for the user to enter a controlled operation mode, which enlarges the risks for the patient and which creates awareness for the user for this condition. | yes | N/A | N/A | N/A
C.2.30 | Does the medical device use an alarm system? | Consider: are the risk of false alarms, missing alarms, disconnected alarm systems, unreliable remote alarm systems, and the medical staff’s possibility of understanding how the alarm system works. Guidance for alarm systems is given in IEC 60601-1-8. | Yes | N/A | N/A | N/A
C.2.31 | In what way(s) might the medical device be deliberately misused? | Consider: are incorrect use of connectors, disabling safety features or alarms, neglect of manufacturer's recommended maintenance. | User error inputting the prescribed settings User Error - ventilator settings inappropriate for patient User Error - not adhering to the PM schedule User Error - O2 not connected or obstructed User Error - cables not connected correctly User Error - Physical damage from impact or vibration, or transport bag fault User Error - environmental temperature or relative humidity exceeded User Error - Physical damage from liquid ingress due to spillage or excessive cleaning liquid | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.32 | Does the medical device hold data critical to patient care? | Consider: the consequence of the data being modified or corrupted. | Yes | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.33 | Is the medical device intended to be mobile or portable? | Consider: necessary grips, handles, wheels, brakes, mechanical stability, and durability. | Portable | Refer to RE00076775 Ventilation Hazards | N/A | N/A
C.2.34 | Does the use of the medical device depend on essential performance? | Consider: for example, the characteristics of the output of life-supporting devices or the operation of an alarm. See IEC60601-1-1 for discussion of performance. | Yes | N/A | N/A | N/A

## Risk Analysis Chart
|  | Document: | 10037183 | Revision: |  |  | W
SECTION 1 | [13 empty] | SECTION 3
RISK ASSESSMENT
HAZARD ANALYSIS | [8 empty] | RISK ANALYSIS (PRE-MITIGATION) |  |  |  |  | RISK CONTROL (POST-MITIGATION) | [7 empty] | VERIFICATION |  | RISK ACCEPTABILITY
Hazard ID | Hazard (may be combined with Hazardous Situation) | Hazardous Situation (may be combined with Hazard) | Potential Cause | Potential Harm Name | Essential Performance (Y/N) | FDA Harms Code | Harm Description | Severity | Req. ID | Prob. of occurrence of Harm | Risk Zone | Occurrence Rationale | Information provided to the User | Risk Control Options Use the following risk control options in the priority order listed |  | Information provided to the User | P1 | P2 | Prob. of occurrence of Harm | Risk Zone | Occurrence Rationale | Verification of Implementation | Verification of Effectiveness | Risk Acceptable? | Risk Acceptance Rationale
[14 empty] | (1st choice) Inherently safe by design | (2nd choice) Protective measures | [5 empty] | P2
VH1 | Alarm - audio volume too high | Patient or user reports alarm volume is too high |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD417 | [6 empty] | Manual - WARNING: Alarm volume should be adjusted with respect to the ventilator’s operating environment and so that the patient's caretakers can hear the alarms. The audible alarm vents located at the front of the device should never be obstructed. The alarmcan be paused with the Alarm Pause function by pressing the ALARM CONTROL key twice once the alarm has been declared. | 2 | 4 | 2 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and no patient harm reported in complaint data. | VHVH41 | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User adjusts alarm volume inappropriately for the environment speaker is out of specification |  | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD32 | [5 empty] | Protective- alarm volume is adjustable 65 to 80 dBA through the preference menu |  | 2 | 2 | 1 | 1
|  |  | DFMEA CPU PCBA: No alarm on reset |  | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD33 | [5 empty] | Background checks for alarm functions |  | 2 | 1 | 1 | 1
[5 empty] | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | PRD296 | [6 empty] | Manual - Operational Verification Checklist - Verify the alarm volume is adapted to the patient environment. | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH2 | Alarm - Delay in response | User or caregiver does not respond to alarm as they do not hear alarm or misinterprets alarm priority | User adjusts alarm volume inappropriately for the environment | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD417 | [6 empty] | Manual - WARNING: Alarm volume should be adjusted with respect to the ventilator’s operating environment and so that the patient's caretakers can hear the alarms. The audible alarm vents located at the front of the device should never be obstructed. The alarm can be paused with the Alarm Pause function by pressing the ALARM CONTROL key twice once the alarm has been declared. | 2 | 3 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefit of using the ventilator for homecare outweighs the potential risk.
|  |  | Use Error - Alarm priority misinterpreted by the user |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD32 | [5 empty] | Protective - alarm volume is adjustable 65 to 80 dBA through the preference menu |  | 2 | 3 | 1 | 1
|  |  | User closes the door to a patient room |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD296 | [6 empty] | Manual - Operational Verification Checklist - Verify the alarm volume is adapted to the patient environment. | 2 | 3 | 1 | 1
|  |  | User doesn't react to all alarms in event of multiple alarm activation |  | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD404 | [6 empty] | Manual - WARNING: When an alarm condition is triggered, or there is evidence of a patient-ventilator fault or problem, examine the patient first before examining the ventilator. | 2 | 2 | 1 | 2
[5 empty] | N |  |  |  | PRD43 |  |  |  |  | Nurse call optional accessory
[5 empty] | Y |  |  |  | PRD162 | [5 empty] | Protective: Visual alarm messages appear on the screen and may flash
VH3 | Alarm Failure | Device or device accessory fails to alarm when it meets alarm condition such as volume, flow, pressure, device alerts. | Total Power Loss | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [6 empty] | Protective- The ventilator has 2 speakers - dual circuits and dual speakers provide back up to alarm function |  | 2 | 3 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA Alarms No detection of High Peak Pressure No detection of Low Pressure Incorrect monitored exhaled tidal volume (VTE)from flow sensor No detection of Apnea No detection of High Breath Rate/Respiratory Rate No detection of Occlusion No detection of "Check Proximal Line 1 " alarm No detection of disconnect No detection of FIO2 alarm condition No main power loss detection |  | Y | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 3 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD50 | [5 empty] | The ventilator performs visual and auditory tests of LEDs and alarm buzzers |  | 2 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | [6 empty] | Protective - each alarm condition has different priority alarms |  | 2 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD404 | [6 empty] | Manual - WARNING: When an alarm condition is triggered, or there is evidence of a patient-ventilator fault or problem, examine the patient first before examining the ventilator. | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD296 PRD437 | [6 empty] | Manual - Operational Verification Checklist If the ventilator fails the alarm tests or if you cannot complete the tests, refer to Chapter 3, “Alarms and Troubleshooting” or call your equipment supplier or Covidien. | 2 | 1 | 1 | 2
|  |  |  | Pneumothorax | N | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 1
[5 empty] | N | 2012 | Symptomatic intervention indicated (e.g. tube placement without sclerosis (no pleurodesis), hospitalization observation, electively place chest tube | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 1
[5 empty] | N |  |  |  | PRD437 | [6 empty] | PB560 - WARNING: If the ventilator fails the alarm tests or if you cannot complete the tests, refer to section 3.8, “Troubleshooting” or call your equipment supplier or Covidien.
[5 empty] | N |  |  |  | PRD436 (560) | [6 empty] | PB560 - WARNING: Verify the functionality of the alarm conditions before connecting the patient to the ventilator
VH4 | False Alarm | Ventilator alarm(s) activated when it should not have occurred | Flow sensor or pressure sensor out of calibration or inaccurate. SW detection criteria is faulty. | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD396 |  |  |  |  | Calibration of the exhalation flow sensor |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA Alarms Incorrect detection of High Peak Pressure Incorrect detection of Low Pressure Incorrect monitored exhaled tidal volume (VTE) from flow sensor Incorrect detection of Apnea Incorrect detection of Occlusion Incorrect detection of "Check Proximal Line 1 " alarm Incorrect detection of disconnect Monitors Incorrect annunciation of FIO2 alarm Incorrect battery level monitor |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD50 | [5 empty] | visual and auditory tests of LEDs and alarm buzzers (POST) |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD357 | [5 empty] | Preventative maintenance schedule - pressure sensor change |  | 2 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD162 | [5 empty] | graphics to display patient response
[5 empty] | N |  |  |  | PRD356 | [6 empty] | WARNING: To ensure proper performance of the ventilator, the preventative maintenance schedule should be followed. For further information contact Covidien.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH5 | No shut down alarm | Audible alarm fails to activate during an unexpected shut down. | User error - lack of service - Buzzer board with battery not serviced | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD50 PRD417 |  |  |  |  | POST |  |  | 2 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA: Security buzzer Buzzer board connector |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [6 empty] | Protective - The ventilator has 2 speakers - dual circuits and dual speakers provide back u to alarm function |  | 2 | 1 | 1 | 1
|  |  | DFMEA Cables - Buzzer board to CPU cable |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD51 |  |  |  |  | Safety net - watchdogs |  |  | 2 | 1 | 1 | 1
|  |  | DFMEA Buzzer PCBA: DC/DC (5V to 12V) Convertor Supply Switch (5V; battery) Power supply loss management Buzzer PCBA back up coin battery Coin Battery Oscillator PWM signal process Buzzers CPU/ Buzzer Board connector Shutdown DC/DC management |  | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD50 | [5 empty] | BUZZER FAULT1 alarm - Defective operation of the buzzers BUZZER FAULT2 alarm - Failure detected in the Very High Priority buzzer BUZZER FAULT3 alarm - Battery Charge Failure due to incorrect voltage BUZZER LOW BATTERY alarm - Buzzer Battery Failure (previously called BUZZER FAULT 4) |  | 2 | 1 | 1 | 2
|  |  | SWFMEA Audio Device Driver |  | N | N/A | Death | 5 | PRD417 | [6 empty] | warning instructing the user to adjust the alarm volume to an appropriate level and to be careful not to obstruct the buzzers outlet. | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD296 | [6 empty] | Manual - Operational Verification Checklist | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
VH6 | False annunciation of security (shut down) alarm | Audible shut down alarm activates during ventilation. | DFMEA - CPU PCBA: Security Buzzer | Anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [5 empty] | Design - Power button and audio pause pressed at the same time will turn the alarm off |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA Buzzer PCBA: Shutdown DC/DC management |  | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD50 | [5 empty] | BUZZER FAULT1 alarm - Defective operation of the buzzers BUZZER FAULT2 alarm - Failure detected in the Very High Priority buzzer BUZZER FAULT3 alarm - Battery Charge Failure due to incorrect voltage BUZZER LOW BATTERY alarm - Buzzer Battery Failure (previously called BUZZER FAULT 4) |  | 2 | 2 | 1 | 1
|  |  | DFMEA Cables - Nurse call system / CPU cables and connectors |  | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD296 | [6 empty] | Manual - Operational Verification Checklist | 2 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD404 | [6 empty] | When an alarm condition is triggered, or there is evidence of a patient-ventilator fault or problem, examine the patient first before examining the ventilator. | 2 | 2 | 1 | 1
VH7 | Alarm - Remote - fail | Remote alarms fail to activate when the ventilator meets an alarm condition. | User Error - Mixing up between open and closed nurse call system | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD41 | [5 empty] | The ventilator connects to either a normally-open nurse call system or a normally closed nurse call system that will transmit any alarm to the nurse call system. | Clinicians Manual - WARNING: To connect the ventilator to a Nurse Call device, contact Covidien to check the ventilator's compatibility with the Nurse Call device and order a suitable connection cable. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User Error - Cable not connected |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD42 | [5 empty] | The ventilator connects to a normally-closed nurse call system that will transmit any alarm to the nurse call system. | Clinicians Manual - WARNING: Do not use Nurse Call devices that operate based on the closure of an electrical circuit, because the devices often do not take into account possible cable disconnection or a total loss of power. Ensure that the Nurse Call device is always connected to the ventilator. | 2 | 2 | 1 | 1
|  |  | DFMEA- CPU PCBA Remote alarm or nurse call Remote alarm connector KEYBOARD MEMBRANE - USER INTERFACE Nurse call / Remote alarm- Connection to ventilator |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD43 | [5 empty] | The ventilator's nurse call is activated when an audible alarm is active and shall be audio paused by pressing the audio paused key of the ventilator. |  | 2 | 2 | 1 | 1
|  |  | Misidentification of nurse call cable |  | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD239 | [6 empty] | Clinicians Manual - WARNING: Before using the Nurse Call system, ensure that its connections are secure and it operates properly. For more information, contact Covidien. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD232 | [6 empty] | The user and clinician manuals include accurate explanations to check the connection of opened and closed nurse call system. | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD267 | [6 empty] | The service manual includes an explanation to check proper connection of the opened and closed nurse call system. Complete a self-test after the cable has been installed and at regular intervals to ensure the system is operating as intended. | 2 | 2 | 1 | 1
[5 empty] | Y | 2484 | Supplemental oxygen is given; COPD patients: Use with caution. Monitored closely. Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | [6 empty] | Protective - primary ventilator alarms are not impacted alarms for volume, flow, pressure, rate, device alerts |  | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Life-threatening consequences; urgent intervention, intubation, or move from non-invasive to invasive ventilatory support indicated. | 4 | [6 empty] | CHECK REMOTE ALARM - No activation of Nurse Call or remote alarm system when an alarm is in progress. |  | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD336 |  |  |  |  | Remote cable compatibilities |  |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH8 | Alarm - Audio - Constant | User is unable to manually reset the audible alarm and/or audible alarm fails to reset when the ventilator alarm violation condition has resolved. | SWFMEA Audio Device Driver | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [5 empty] | Design Power cycling resets the ventilator clearing the issue |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD162 | [5 empty] | graphics to display patient response |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD50 | [5 empty] | POST - performs visual and auditory tests of LEDs and alarm buzzers |  | 1 | 2 | 1 | 1
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 1 | 2 | 1 | 1
VH9 | Alarm - Visual - Constant | User is unable to manually reset the visual alarm and/or visual alarm fails to reset when the ventilator alarm violation condition has resolved. | DFMEA- CPU PCBA keyboard connector | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | PRD32 PRD33 | [5 empty] | Audible alarms annunciate |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD162 | [5 empty] | graphics to display patient response |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 1 | 2 | 1 | 1
VH10 | Alarm - Audio - fails to enunciate/sound | Audio alarms fail to activate when the ventilator meets the alarm violation conditions. | DFMEA CPU PCBA Watchdog | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [6 empty] | Protective - The ventilator has 2 speakers - dual circuits and dual speakers provide back u to alarm function | Operational Verification Test | 2 | 3 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA Audio Device Driver |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD51 | [5 empty] | Safety net -Alarm system that continuously monitors both patient and machine. The machine-related alarm conditions are factory set, |  | 2 | 3 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1
|  |  |  | Respiratory failure | Y | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | [6 empty] | alarms for volume, flow, pressure, rate, device alerts |  | 2 | 1 | 1 | 1
[5 empty] | Y | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD404 | [6 empty] | When an alarm condition is triggered, or there is evidence of a patient-ventilator fault or problem, examine the patient first before examining the ventilator. | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 1 | 1 | 2
|  |  |  | Pneumothorax | N | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 1
[5 empty] | N | 2012 | Small to Moderate Pneumothorax, electively place chest tube or observation | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 1
[6 empty] | 2012 | Ventilation: Pneumothorax requiring placement chest tube with hospitalization or prolonged hospitalization | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 2012 | Life-threatening consequences; Tension pneumothorax; airway compromise requiring intubation; urgent surgical intervention or emergency Chest tube placement indicated | 4 | [8 empty] | 2 | 1 | 1 | 2
VH11 | Alarm - Visual - fails to illuminate/light | Visual alarm indicator and/or alarm message fail to activate when the ventilator meets the alarm violation conditions | DFMEA CPU PCBA - Power Supply Board connector - keyboard connector | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD296 | [6 empty] | Manual - Operational Verification Checklist | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA KEYBOARD MEMBRANE USER INTERFACE |  | Y | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD33 PRD34 | [5 empty] | audio alarm will annunciate |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD162 | [5 empty] | graphics to display patient response |  | 1 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH12 | Alarm - Volume too low | Patient or user reports alarm volume is too low. | User adjusts alarm volume inappropriately for the environment | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD32 |  |  |  |  | Graphics display includes a Parameter adjustment bar for alarm volume | operational verification checklist | Labelling - Manual - WARNING: Alarm volume should be adjusted with respect to the ventilator’s operating environment and so that the patient's caretakers can hear the alarms. The audible alarm vents located at the front of the device should never be obstructed. The alarm can be paused with the Alarm Pause function by pressing the ALARM CONTROL key twice once the alarm has been declared. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Obstruction of the sound diffusion holes |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [6 empty] | Protective - alarm volume is adjustable 65 to 80 dBA |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD32 PRD33 | [5 empty] | Audible alarms annunciate |  | 2 | 3 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD34 | [5 empty] | Buzzer alarm |  | 2 | 3 | 1 | 2
[5 empty] | N | N/A | Death | 5 | [6 empty] | Visual alarms LEDs |  | 2 | 3 | 1 | 2
|  |  |  | Respiratory failure | Y | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD417 | [5 empty] | Alarm message on display |  | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD435 | [6 empty] | warning instructing the user to adjust the alarm volume to an appropriate level and to be careful not to obstruct the buzzers outlet. | 2 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD433 (560) | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | [7 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
VH13 | Deleted
VH14 | Battery fails to charge | Battery fails to charge (internal or external). | User error - Ventilator has not been charged whilst in storage | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD213 |  |  |  |  | Whenever an AC power supply is available, the ventilator recharges the internal battery |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA: Power Supply PCBA - Current control for battery charge (Constant current Generator) ACDC POWER cables ports |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD217 |  |  |  |  | The ventilator is able to estimate battery ageing. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD220 |  |  |  |  | When the Ventilator is in Active Ventilation State, the time to fully charge the internal battery from a complete discharge does not exceed 13 hours. | Battery charge level is displayed in time whilst ventilating and in % while on standby |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD157 |  |  |  |  | Real time battery life indicator will adjust the indication of battery time remaining depending on current ventilator settings and/or setting changes |  | Manual - WARNING: Before using the ventilator’s internal battery, ensure that the battery is fully charged and that the charge holds. Back up ventilators or those in storage should be connected to an AC power source to protect the integrity of the battery. | 2 | 1 | 1 | 2
[5 empty] | Y | N/A | Death | 5 | PRD199 | [5 empty] | Battery alarms EMPTY BATTERY ALARM - Internal battery capacity < 10 min. or 3% CHECK BATTERY CHARGE ALARM - Internal battery charging failure |  | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD243 | [6 empty] | Manual - WARNING: Before using the ventilator’s internal battery, ensure that the battery is fully charged and that the charge holds. Back up ventilators or those in storage should be connected to an AC power source to protect the integrity of the battery. | 2 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD423 | [6 empty] | WARNING: The maximum recommended shelf life of the internal battery is two (2) years. Do not use a battery that has been stored for two years prior to its first use. | 2 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD377 |  |  |  |  | The ventilator packaging includes the storage duration recommended before recharging the internal batteries. |  |  | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD155 |  |  |  |  | Design - Battery Power indicator flashes when battery is charging |  |  | 2 | 1 | 1 | 2
[5 empty] | Y |  |  |  | PRD309 | [5 empty] | LOW BATTERY alarm | Manual - WARNING: When the “LOW BATTERY“ alarm is triggered, immediately connect the ventilator to an AC power supply to maintain ventilation and recharge the internal battery.
[5 empty] | N | [10 empty] | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients.
[5 empty] | Y |  |  |  | PRD197 | [5 empty] | AC POWER DISCONNECTION alarm
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD216 | [6 empty] | PB560 - WARNING: Even if the “INTERNAL BATTERY” charging indicator is off, charging of the battery may sometimes be incomplete if the ambient temperature is above 40°C (104°F) because of the battery’s internal heat safety device.
VH15 | Battery level indication is inaccurate | Displayed Battery level indicator displays incorrect charge status of battery. | SWFMEA: - Incorrect battery level monitor Software algorithms are showing incorrect % | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD155 |  |  |  |  | Design - Battery Power indicator flashes when battery is charging |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | PM schedule not adhered to |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD156 |  |  |  |  | Real time battery life indicator will adjust the indication of battery percent available depending on current ventilator settings and/or setting changes |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD157 |  |  |  |  | Real time battery life indicator will adjust the indication of battery time remaining depending on current ventilator settings and/or setting changes |  | Manual - WARNING: Before using the ventilator’s internal battery, ensure that the battery is fully charged and that the charge holds. Back up ventilators or those in storage should be connected to an AC power source to protect the integrity of the battery. | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD158 |  |  |  |  | The ventilator provides a visual indication of the active ventilator power source (AC, DC or Battery) that is operational in all ventilator states and modes. |  | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD219 |  |  |  |  | When the Ventilator is in Ventilation Stand-by State, the time to fully charge the internal battery from a complete discharge does not exceed 6 hours. |  |  | 1 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD220 |  |  |  |  | When the Ventilator is in Active Ventilation State, the time to fully charge the internal battery from a complete discharge does not exceed 13 hours. | Battery charge level is displayed in time whilst ventilating and in % while on standby |  | 1 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD357 |  |  |  |  | PM schedule - battery charge level information is tested during PM |  |  | 1 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD243 | [6 empty] | Manual - WARNING: Before using the ventilator’s internal battery, ensure that the battery is fully charged and that the charge holds. Back up ventilators or those in storage should be connected to an AC power source to protect the integrity of the battery. | 1 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD423 | [6 empty] | WARNING: The maximum recommended shelf life of the internal battery is two (2) years. Do not use a battery that has been stored for two years prior to its first use. WARNING: Periodic recharging is important to help maximize useful life of the battery. Do not store the internal battery for extended periods, without recharging, as this may reduce the maximum life. | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD377 |  |  |  |  | The ventilator packaging shall include the storage duration recommended before recharging the internal batteries.
[5 empty] | Y |  |  |  | PRD199 | [5 empty] | Battery alarm CHECK BATTERY CHARGE ALARM - Internal battery charging failure
[5 empty] | Y |  |  |  | PRD194 | [5 empty] | Low Battery Alarm: A ventilator operating on internal battery shall provide a high priority alarm when the estimated remaining operating time is 30 minutes or below.
[5 empty] | N |  |  |  | PRD195 | [5 empty] | Empty Battery Alarm: A ventilator operating on internal battery shall provide a high priority alarm when the estimated remaining operating time is 10 minutes or below.
[5 empty] | Y |  |  |  | PRD309 | [5 empty] | LOW BATTERY alarm | Manual - WARNING: When the “LOW BATTERY“ alarm is triggered, immediately connect the ventilator to an AC power supply to maintain ventilation and recharge the internal battery.
[5 empty] | Y |  |  |  | PRD197 | [5 empty] | AC POWER DISCONNECTION alarm
[5 empty] | Y |  |  |  | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss
[5 empty] | N | [10 empty] | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH16 | Battery does not power ventilator | No power when on battery. Ventilator fails to switch to battery. | User does not connect to AC or DC power No Preventative Maintenance - battery life exceeded | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD158 |  |  |  |  | Visual LED indication of the active ventilator power source (AC, DC or Battery) that is operational in all ventilator states and modes. | - |  | 1 | 3 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA: Power Supply PCBA - reference 3 V supply Micro-controller Cooling fan control (ventilator) Cooling fan connector |  | Y | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD197 | [5 empty] | AC POWER DISCONNECTION alarm |  | 1 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD198 | [5 empty] | DC POWER DISCONNECTION alarm |  | 1 | 1 | 1 | 1
[5 empty] | Y | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD199 | [5 empty] | Battery alarms EMPTY BATTERY ALARM - Internal battery capacity < 10 min. or 3% CHECK BATTERY CHARGE ALARM - Internal battery charging failure |  | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Respiratory failure |  | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD423 | [7 empty] | 1 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD407 | [6 empty] | WARNING: Ensure that the ventilator’s internal battery is fully charged before connecting the ventilator to an external DC power source. Powering the ventilator using an external 12– 30 VDC power source (via the DC power cable) does not enable its internal battery to recharge. | 1 | 1 | 1 | 2
|  |  |  | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD212 | [6 empty] | PB560 WARNING: Due to its limited internal battery’s reserve capacity, the ventilator should only be operated on the internal battery when no other power source is available. Ensure that the internal battery never becomes fully discharged. | 1 | 3 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | [8 empty] | 1 | 3 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD435 | [7 empty] | 1 | 1 | 1 | 2
[6 empty] | 1916 | Death | 5 | PRD433 (560) | [7 empty] | 1 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) |  | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2477 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH17 | Battery is unable to be removed | Battery is stuck or jammed in place. |  | This hazard is not applicable to this ventilator - This failure has been designed out as the PB500 series battery is changed during service. All service engineers must ne trained. Service departments should have the tools necessary to unscrew the door and to drill out the screws if the screws are damaged.
VH18 | Biohazard Exposure | During the use or handling of the device or system (e.g. device surface, circuit, humidifiers) user or patient is exposed to pathogens. | Re-use of single patient use components | Infection | N | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | PRD259 | [6 empty] | WARNING: Single Use accessories should not be reused | 2 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | contaminated component package Inadequate package design Improper package handling |  | N | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth or inhalation, Incision/drainage see abscess | 2 | PRD260 | [6 empty] | WARNING: Before opening the packaging for the Patient Circuit, ensure that no damage is evident to the packaging or its contents. Do not use if evidence of damage exists. | 2 | 1 | 1 | 1
|  |  | User Error - Water trap condensate spilled on removal |  | N | 1930 | Symptomatic; IV antibiotics required with increased ventilatory support. An infection that requires more advanced medical intervention | 3 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 1 | 1 | 1
|  |  | Contamination of blower vanes due to air inlet filter not being in place |  | N | 1930 | Symptomatic; IV antibiotics required, increased ventilatory support, isolation, refractory to treatment. | 4 | PRD35 | [5 empty] | The air inlet filter is designed such that an unclean state can be easily detected |  | 2 | 1 | 1 | 2
|  |  | DFMEA FILTRATION SYSTEM - Exhalation Bacteria filter Inspiratory Bacteria Filter |  | N | 1930 | Death | 5 | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator. | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 1
|  |  |  | Cross Patient Exposure To Body Fluids | N | 1745 | No medical treatment required if in contact with in-tack skin | 1 | PRD280 |  |  |  |  | The ventilator packaging and shipping complies with ADR and IATA regulations depending on the level of dangerous goods contained in the device. All finished goods units deemed critical to the safe operation of the ventilator will be subjected to the ISTA-2A test standard ensuring that product integrity is maintained after the distribution environment. |  |  | 2 | 1 | 1 | 1
[5 empty] | N |  | prophylactic treatment required for exposure to blood or body fluid | 2 | [7 empty] | Manual - WARNING: The exhalation block is intended for single use by a single patient . It may periodically be cleaned, but it cannot be disinfected or sterilized. To maintain good measurement quality when used continuously, clean the exhalation block periodically. The exhalation block should be changed every 4 months and cannot be reused with any other patient. | 2 | 1 | 1 | 1
[5 empty] | N |  | extended prophylactic treatment required for exposure to blood or body fluid | 3 | PRD405 | [6 empty] | WARNING: A patient treated by mechanical ventilation is highly vulnerable to the risks of infection. Dirty or contaminated equipment is a potential source of infection. Clean the ventilator and its accessories regularly and systematically before and after each use and following any maintenance procedure to reduce the risks of infection. The use of a bacterial filter at the ventilator’s outlet – or both ports if a double-limb circuit is used – is highly recommended. Refer to chapter 7, “Cleaning”. To reduce the risk of infection, wash your hands thoroughly before and after handling the ventilator or its accessories. | 2 | 1 | 1 | 1
[5 empty] | N |  | life threatening infection, hospitalization or prolonged hospitalization, or greater than 3 drugs | 4 | [5 empty] | DAR filters labelling indicates filtration efficiencies | Acceptable DAR filters are listed as accessories in the User Manual |  | 2 | 1 | 1 | 2
[5 empty] | N |  | Death | 5 | [6 empty] | Air Inlet filter is on the Preventative Maintenance schedule |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD231 | [6 empty] | WARNING: The patient circuit is intended for single use by a single patient and should be changed according to the manufacturer’s recommendations and according to the patient circuit lifetime. Refer to the instructions for use supplied by the manufacturer of the patient circuit (included with the ventilator) and Chapter 4, “Installation and Assembly.”
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD1 | [6 empty] | Warning: Lay responsible organisation (e.g. Homecare Vent Provider) must contact local authorities to determine the proper method of disposal of potentially bio hazardous parts and accessories.
VH19 | Breathing circuit leak | Breathing circuit has a leak | Use of a single limb patient circuit with a vented NIV interface and with a volume mode leading to insufficient ventilation | Hypoventilation | Y | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD350 | [5 empty] | Low pressure alarms CHECK EXH VALVE CHECK PROXIMAL LINE1 NO PROXIMAL LINE2 |  | 2 | 3 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Leak in patient circuit |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 3 | 1 | 1
|  |  | Patient movement, Patient circuit accidentally disconnected (ventilation loss) |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD273 | [5 empty] | All device operational connectors and ports are labeled |  | 2 | 2 | 1 | 1
|  |  | DFMEA humidifier - connection to the breathing circuit water traps - connection to the breathing circuit |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD248 | [6 empty] | Manual WARNING: The patient circuit should always be positioned to avoid hindering the patient's movements, to prevent accidental disconnection or leakage, and to minimize the risk of patient strangulation. | 2 | 1 | 1 | 2
|  |  | Use Error - Incorrect set-up of patient circuit |  | Y | 1916 | Death | 5 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | Y | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD 180 PRD181 PRD188 | [5 empty] | HIGH VTI ALARM LOW VTI Alarm |  | 2 | 3 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD184 | [5 empty] | NO PROXIMAL LINE2 alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD320 | [5 empty] | Leakage alarms HIGH LEAKAGE Alarm EXH VALVE LEAKAGE ALARM E SENS FAULT OR CIRC LEAK |  | 2 | 1 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD189 (560) | [5 empty] | HIGH RATE alarm (due to autotriggering) |  | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD240 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. On a DAILY basis, inspect the patient circuit to ensure that it shows no signs of damage, is properly connected, and is operating correctly without leakage. | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD306 | [6 empty] | Manual - WARNING: The patient circuit should not be changed during ventilation. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD230 | [6 empty] | Description of patient circuit | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD415 | [6 empty] | Manual - WARNING: When using non-invasive ventilation (NIV), without an exhalation valve, use a vented nose or face mask or a non vented combined with a leak accessory. When using noninvasive ventilation (NIV), with an exhalation valve, use a non vented mask. | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD245 | [6 empty] | WARNING: The patient circuit is intended for single use by a single patient and should be changed according to the manufacturer’s recommendations and according to the patient circuit lifetime. | 2 | 1 | 1 | 2
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD280 |  |  |  |  | Circuit packaging is ISTA-2 complaint |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [5 empty] | Design - Water traps are screwed in and keyed |  |  | 2 | 1 | 1 | 1
[5 empty] | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [7 empty] | WARNING: Users must always possess an additional breathing circuit and exhalation valve | 2 | 1 | 1 | 1
[5 empty] | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [7 empty] | The circuit check menu displays instructions for the steps needed to conduct the test and a confirmation step and a warning to ensure the patient is not connected prior to conducting the test. | 2 | 1 | 1 | 2
[5 empty] | N | 2328 | See cause of anxiety | 5 | [5 empty] | Design - Hose connections are designed that the fittings are not easily pulled out, optional swivel adapter for patient wye |  |  | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, Pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD459 |  |  |  |  | Circuit leak test |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD411 | [6 empty] | WARNING: The level of inspiratory resistance of the circuit and accessories (bacteria filter, humidifier, and so on) must be as low as possible. Settings – particularly the PATIENT DISCONNECTION alarm, High inspired volume (High VTI), and Low inspired volume (Low VTI) settings – must be periodically adjusted according to changes in the patient circuit resistance – especially when filters are replaced. | 2 | 1 | 1 | 2
VH20 | Breathing circuit occluded, pinched or kinked | Patient Breathing circuit has a restriction causing decreased pressure, flow, or volume. | Mucus impaction, not sufficient ventilation. | Hypoventilation | Y | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm | WARNING: The patient circuit should always be positioned to avoid hindering the patient's movements, to prevent accidental disconnection or leakage, and to minimize the risk of patient strangulation. | 2 | 3 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User error - water inside exhalation block due to humidifier use |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms |  | 2 | 2 | 1 | 1
|  |  | Use error - Patient circuit occluded, pinched or kinked |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD189 (560) | [5 empty] | CHECK EXHALATION VALVE alarm (PB560) |  | 2 | 1 | 1 | 1
|  |  | Patient circuit arm faulty or not used properly installed on the cart. |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD174 | [5 empty] | Low pressure alarms CHECK EXH VALVE CHECK PROXIMAL LINE1 |  | 2 | 1 | 1 | 2
|  |  | SWFMEA No detection of Occlusion |  | N | 1916 | Death | 5 | PRD306 | [6 empty] | Manual - WARNING: The patient circuit should not be changed during ventilation. | 2 | 1 | 1 | 2
|  |  |  | Lung, Over inflation | N | 2397 | Asymptomatic-Observed | 1 | [7 empty] | WARNING: Users must always possess an additional breathing circuit and exhalation valve | 2 | 2 | 1 | 1
[5 empty] | N | 2397 | Small ↑ mean airway pressure, requiring minor change in ventilator settings. | 2 | PRD230 | [6 empty] | Description of patient circuit | 2 | 1 | 1 | 1
[5 empty] | N | 2397 | Moderate ↑ mean airway pressure, radiographical signs, requiring major change in ventilator settings. | 3 | PRD234 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. | 2 | 1 | 1 | 1
[5 empty] | N | 2397 | Critical↑ mean airway pressure, radiographical signs parenchemal destruction (e.g. blebs), requiring major change in ventilator therapy (e.g., ECMO, HFV) | 4 | PRD248 | [6 empty] | Manual - WARNING: The patient circuit should always be positioned to avoid hindering the patient's movements, to prevent accidental disconnection or leakage, and to minimize the risk of patient strangulation. | 2 | 1 | 1 | 2
[5 empty] | N | 2397 | Death | 5 | PRD273 | [5 empty] | All device operational connectors and ports are labeled |  | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD412 | [6 empty] | WARNING: Always position a humidification device so that it is lower than both the ventilator and the patient. Use water traps, if necessary, to limit water in the patient circuit and precautions should be taken when discarding the fluid in the water trap, discard per local ordinance for proper disposal. If a heated humidifier is used, you should always monitor the temperature of the gas delivered to the patient. Gas delivered from the ventilator that becomes too hot may burn the patient's airway. | 2 | 2 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 1 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm |  | 2 | 1 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD 180 PRD181 PRD188 | [5 empty] | HIGH VTI ALARM LOW VTI Alarm |  | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, Pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm |  | 2 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD173 | [5 empty] | HIGH RATE alarm |  | 2 | 1 | 1 | 2
[5 empty] | Y |  |  |  | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD411 | [6 empty] | WARNING: Resistance of the exhalation valve and accessories (water traps, filters, HMEs, etc.) must be as low as possible. WARNING: The exhalation valve must allow rapid discharge of the circuit pressure. Ensure that the exhalation valve is always clean and its evacuation aperture (exhaust port) is never obstructed.
VH21 | Breath not delivered | Ventilator does not deliver breaths per prescribed setting | User Error - Circuit Check is enabled while the patient is being ventilated or connected to the ventilator. | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD459 |  |  |  |  | Circuit leak test |  |  | 2 | 3 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User Error - ventilator settings inappropriate for patient (e.g. I sense, E sense) |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD460 PRD461 |  |  |  |  | The circuit check menu displays instructions for the steps needed to conduct the test and a confirmation step and a warning to ensure the patient is not connected prior to conducting the test. |  |  | 2 | 2 | 1 | 1
|  |  | DFMEA Pneumatic fittings - Inspiratory flow sensor fitting |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 1 | 1 | 1
|  |  | SWFMEA Breath Delivery Initiates breaths incorrectly Alarms No detection of Apnea No detection of disconnect |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD237 | [6 empty] | WARNING: Before operating the ventilator, read, understand, and strictly follow the information contained in Chapter 1, “Safety Information.” | 2 | 1 | 1 | 2
|  |  | Use Error - Incorrect set-up of patient circuit |  | N | 1916 | Death | 5 | PRD413 | [6 empty] | WARNING: Adding accessories to the ventilator breathing circuit, such as a humidifier and water trap(s), may result in a decrease in tidal volume delivered to the patient due to the added compressible volume of the accessory. Always assure that the patient is receiving the appropriate inspired volume when altering the breathing circuit configuration. | 2 | 1 | 1 | 2
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 2 | 3 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD289 | [5 empty] | Leak is displayed |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD320 | [5 empty] | Leakage alarms HIGH LEAKAGE Alarm EXH VALVE LEAKAGE ALARM E SENS FAULT OR CIRC LEAK |  | 2 | 1 | 1 | 1
[5 empty] | Y | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | PRD172 | [5 empty] | Apnea alarms |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD173 | [5 empty] | HIGH RATE alarm
[5 empty] | N |  |  |  | PRD74 |  |  |  |  | PEEP adjustment range
[5 empty] | N |  |  |  | PRD102 |  |  |  |  | APNEA TIME adjustment range
[5 empty] | N |  |  |  | PRD388 PRD87 |  |  |  |  | Inspiratory Sensitivity settings Exhalation sensitivity settings
[5 empty] | Y |  |  |  | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms
[5 empty] | Y |  |  |  | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm
[5 empty] | N |  |  |  | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. Provide the patient with an alternate means of ventilation before conducting any verification tests.
[5 empty] | N |  |  |  | PRD74 | [6 empty] | PEEP adjustment range
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD54 | [6 empty] | PB560 - WARNING: A continuous alarm condition will be activated if the ventilator power switch is turned off while ventilation is in progress. When the power switch is turned back on again, ventilation will resume without having to press the VENTILATION ON/OFF button.
[5 empty] | N |  |  |  | PRD432 (560) | [6 empty] | PB560 - WARNING: Do not allow a patient to remain connected to the ventilator when ventilation is stopped, because a substantial quantity of exhalation gas, primarily carbon dioxide, may be inhaled by the patient. In some circumstances, inhaling carbon dioxide may lead to under-ventilation, suffocation, and serious injury or death.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | WARNING: Always have immediate access to an alternative means of ventilation, which is ready for use, to avoid patient death or serious injury
[5 empty] | N |  |  |  | PRD1 | [6 empty] | WARNING: When using the ventilator in a carrying case, only use a carrying case that is listed in the instructions for use to prevent adverse ventilator performance which can consequently result in Patient death
[5 empty] | N |  |  |  | PRD1 | [6 empty] | WARNING: Do not cover the ventilator or place in a position that affects proper operation”, eg. front or lateral opening.
[5 empty] | N |  |  |  | PRD320a | [6 empty] | WARNING: The setting of the Max Leak alarm must be adjusted for the patient, but must also be set low enough to allow the HIGH LEAKAGE alarm to trigger properly. Perform the Max Leak Test (refer to section E.7, “Max Leak Test”) to ensure the Max Leak alarm is functionally properly . This alarm only applies to Leak configuration (NIV)
[5 empty] | N |  |  |  | PRD424a | [6 empty] | WARNING: Ensure that the ventilator is switched off and disconnected from all external power supplies before installation.
[5 empty] | N |  |  |  | PRD320a | [6 empty] | WARNING: The Max Leak alarm setting must be adjusted for the patient, but must also be set low enough to allow the High Leakage alarm to trigger properly. Perform the max leak test (see section F.2, Max Leak Test (Only NIV)) to ensure that the alarm is functioning properly. This alarm only applies to leak configuration (NIV).
VH22 | Ventilator autotrigger | Inappropriate triggering of ventilation when the patient is not attempting to initiate a breath(e.g. inappropriate sensitivity, cardiogenic noise, leaks) | User Error - Incorrect set up by Clinician - too sensitive settings | Lung, Over inflation | N | 2397 | Asymptomatic-Observed | 1 | PRD161 |  |  |  |  | Indication of patient triggered breath |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | leaks |  | Y | 2397 | Small ↑ mean airway pressure, requiring minor change in ventilator settings. | 2 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 2 | 2 | 1 | 1
|  |  | DFMEA CPU PCBA - Inspiratory flow measure PNEUMATIC FITTINGS - flow sensor fitting |  | Y | 2397 | Moderate ↑ mean airway pressure, radiographical signs, requiring major change in ventilator settings. | 3 | PRD320 | [5 empty] | Leakage alarms HIGH LEAKAGE Alarm EXH VALVE LEAKAGE ALARM E SENS FAULT OR CIRC LEAK |  | 2 | 1 | 1 | 1
|  |  | SWFMEA Breath delivery Initiates breaths incorrectly Alarms No detection of High Breath Rate/Respiratory Rate |  | N | 2397 | Critical↑ mean airway pressure, radiographical signs parenchemal destruction (e.g. blebs), requiring major change in ventilator therapy (e.g., ECMO, HFV) | 4 | PRD173 | [5 empty] | HIGH RATE alarm |  | 2 | 1 | 1 | 2
[5 empty] | N | 2397 | Death | 5 | PRD191 | [5 empty] | BREATH TIME CYCLED Alarm |  | 2 | 1 | 1 | 2
|  |  |  | Hyperventilation | N | 1910 | Asymptomatic; PaCO2 or ETCO2 low. Mild tachypnea. | 1 | PRD359 |  |  |  |  | CHECKLIST |  | WARNING: Ensure that the I Sens setting is not set to OFF when ventilating patients capable of triggering spontaneous breaths. Carefully modify the trigger threshold setting to reduce the risk of ventilator autotriggering. Level 0P, the most sensitive inspiratory trigger, is recommended for pediatric use. For an adult, this setting may result in ventilator autotriggering. | 2 | 3 | 1 | 1
[5 empty] | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD307 | [6 empty] | WARNING: If APNEA TIME is set to a value higher than 60/Control R then the APNEA alarm will not activate. If an APNEA alarm is required, set the APNEA setting to YES in the Preferences Menu. | 2 | 2 | 1 | 1
[5 empty] | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD447 (560) | [6 empty] | WARNING: The Apnea Alarm should be set to YES for ventilator dependent patients. | 2 | 1 | 1 | 1
|  |  |  | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD296 | [6 empty] | Manual - Operational Verification Checklist - Ensure the patient breathing circuit is correctly attached to the ventilator | 2 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 2 | 1 | 1
[5 empty] | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm |  | 2 | 1 | 1 | 1
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 3 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD86 |  |  |  |  | EXH SENS setting |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD83 |  |  |  |  | adjustable trigger sensitivity setting |  |  | 2 | 1 | 1 | 1
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH23 | Burning smell or odor | Device or device accessory exhibits a burning smell or odor |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [5 empty] | Design - Cooling fan will withdraw any smoke from the interior of the ventilator |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA Blower Turbine |  | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [5 empty] | Design - No inlet into the gas pathway from the ventilator electrical interior |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [5 empty] | All PCBAs are IPC compliant which includes flame retardant materials |  |  | 2 | 1 | 1 | 1
[5 empty] | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
|  |  |  | Chemical (gas or liquid)exposure , Hazardous | N | 2570-Chemical Exposure | Minor Inflammation, no treatment needed | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N |  | Temporary injury or impairment that requires minor professional medical intervention | 2 | PRD1 |  |  |  |  | PB560 Ventilator is designed to meet the The ventilator complies with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1 |  |  | 2 | 2 | 1 | 1
[7 empty] | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | [8 empty] | 2 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 3 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 3 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
VH24 | Cable Entanglement | Device or device accessory that becomes entangled around patient or user | Patient circuit wraps around neck, legs, arms - patient or user harm | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | [6 empty] | Breathing circuits come equipped with cable management clips. |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Patient circuit wraps around neck, legs, arms - ventilator disconnect |  | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. | 2 | PRD248 | [6 empty] | WARNING: The patient circuit should always be positioned to avoid hindering the patient's movements, to prevent accidental disconnection or leakage, and to minimize the risk of patient strangulation. | 1 | 1 | 1 | 1
|  |  | Power cord wraps around neck, legs, arms - patient harm |  | N | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue | 3 | PRD428 | [6 empty] | WARNING: Do not leave power cables lying on the ground where they may pose a hazard. | 1 | 1 | 1 | 1
|  |  | Power cord wraps around neck, legs, arms - ventilator disconnect |  | N | 2104 | Sudden significant injury to tissue that requires immediate attention to prevent a catastrophic event | 4 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 1 | 1 | 1 | 2
|  |  | SWFMEA Inaccurate calibration of the ventilator when performing Performance Verification Tests. Inaccurate settings change when performing Performance Verification Tests. |  | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH25 | Calibration fail | Device did not pass calibration | DFMEA CPU PCBA Inspiratory flow measure | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD365 |  |  |  |  | Patient circuit type may be selected prior to calibration |  |  | 3 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | PM schedule not adhered to |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [6 empty] | CALIBRATION FAIL alarm |  | 3 | 1 | 1 | 1
|  |  | SWFMEA PB500 VTS Inaccurate calibration of the ventilator when performing Performance Verification Tests. Inaccurate settings change when performing Performance Verification Tests. |  | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD321 (560) PRD322 (560) | [5 empty] | PB560 CHECK FIO2 SENSOR alarm LOW FIO2 alarm HIGH FIO2 alarm FIO2 SENSOR MISSING alarm CALIBRATE FIO2 alarm | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FIO2 kit) that features a minimum and maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 3 | 1 | 1 | 1
|  |  | Use Error - inaccurate calibration | Oxygen saturation, Low (Hypoxemia) | Y | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm |  | 3 | 1 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD 180 PRD181 PRD188 | [5 empty] | HIGH VTI ALARM LOW VTI Alarm |  | 3 | 1 | 1 | 1
|  |  |  | Hypoventilation (COPD) | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD357 | [5 empty] | Preventative maintenance schedule - pressure sensor change |  | 3 | 1 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | [7 empty] | Note: Calibration may be done with either an adult or pediatric circuit; however, Pediatric YES/NO must be appropriately selected in the ventilator Preferences Menu. | 3 | 1 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | [7 empty] | Each time the exhalation block or circuit is removed and reinstalled or after installing a new exhalation block, the exhalation flow sensor must be recalibrated before the using the ventilator. This process is automatic and does not require the use of a measurement device. | 3 | 1 | 1 | 1
[5 empty] | N | [10 empty] | WARNING: When an exhalation block is set up, each time it is removed, or after installing a new exhalation block on the machine, it is essential that the exhalation flow sensor be recalibrated before the exhalation block is used. Refer to section 10.1, “Calibrating the Exhalation Flow Sensor”.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH26 | Chemical Toxicity | Exposure of chemicals or cleaning agents to patient, user, or environment | Inappropriate material selection for the gas pathway | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PRD17 |  |  |  |  | gas pathway materials biocompatibility applying testing requirements per Environmental Protection Agency (EPA) method TO-15 |  |  | 2 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | breathing circuit that has not been tested for biocompatibility |  | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | PRD261 | [6 empty] | WARNING: Never use a liquid cleaner inside the patient circuit, or on any component of a gas pathway. Clean the patient circuit only as specified by the manufacturer's instructions. | 2 | 1 | 1 | 1
|  |  | inappropriate cleaning chemicals - breathing circuit |  | N | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | PRD47 |  |  |  |  | All recommended patient circuits shall comply with ISO 5367. |  |  | 2 | 1 | 1 | 1
|  |  | inappropriate cleaning chemicals - external surfaces of ventilator | Chemical exposure, Hazardous | N | 2570 | Minor Inflammation, no treatment needed | 1 | PRD24 | [6 empty] | cleaning instruction include appropriate methods and chemicals | 2 | 1 | 1 | 1
|  |  | Harmful substances leaking from internal battery case |  | N | 2570 | Temporary injury or impairment that requires minor professional medical intervention | 2 | PRD374 |  |  |  |  | The ventilator complies with directive 2006/66/EC on batteries and accumulators and waste batteries and accumulators, 2006. |  | WARNING: For environmental protection, the ventilator and its components, whatever their respective conditions of operation, cannot be disposed of with household waste and must be submitted for suitable selective collection and possible recycling. Observe all applicable regulations when disposing of the ventilator and any of its components. | 2 | 1 | 1 | 1
|  |  | External battery leakage |  | N | 2570 | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | PRD375 |  |  |  |  | The ventilator complies with directive 2002/96/EC, 2003 on waste electrical and electronic equipment (WEEE), 2003. |  | Warnings Regarding the Environment • For environmental protection, the ventilator and its components, whatever their respective conditions of operation, cannot be disposed of with household waste and must be submitted for suitable selective collection and possible recycling. Observe all applicable regulations when disposing of the ventilator and any of its components. • Follow local governing ordinances and recycling regulations regarding disposal or recycling of batteries and other device components. | 2 | 1 | 1 | 1
|  |  | inappropriate cleaning chemicals - breathing circuit |  | N |  |  |  | PRD373 PRD265 PRD266 PRD270 | [5 empty] | The ventilator and battery are labeled to indicate that the device shall be considered as waste electrical and electronic equipment. | Recycle electrical and electronic items with this marking according to local regulations (EN50419). Do not discard with ordinary household waste. This includes the ventilator and the battery. The battery is labeled to indicate that the battery shall be considered as waste electrical and electronic equipment.
|  |  | Parts containing phthalates |  | N |  |  |  | PRD302 PRD360 | [5 empty] | The transport bag (dual bag) is labeled with washing machine cleaning instructions | WARNING: Regularly clean the ventilator’s Dual Bag according to manufacturer’s recommendations.
[5 empty] | N |  |  |  | PRD405 | [5 empty] | Instruction to regularly clean the ventilator
[5 empty] | N |  |  |  | PRD427 | [5 empty] | The exterior of the ventilator is cleanable with a cloth or sponge lightly moistened with a bactericide or germicide solution | WARNING: Use all cleaning solutions and products with caution. Read and follow the instructions associated with the cleaning solutions you use to clean your ventilator. Use only those solutions listed in Table 7-1.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD464 |  |  |  |  | The ventilator complies with Biological Evaluation of medical devices Part 1: Evaluation and testing within a risk management process (EN ISO 10993-1).
VH27 | Compressor Inoperative | Inability of the compressor to deliver gas for Breath Delivery |  | N/A for this ventilator as there is no external compressor
VH28 | Contamination | Introduction of debris, dust, or contaminants to patient, user, or environment. | Foreign material in Lungs - Inlet filter is missing | Infection | N | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | PRD17 |  |  |  |  | The gas pathway materials biocompatibility applying testing requirements per Environmental Protection Agency (EPA) method TO-15 | - |  | 2 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Dirt gets into lung - Inlet filter very dirty - PM schedule has not been followed |  | N | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth or inhalation, Incision/drainage see abscess | 2 | PRD35 | [5 empty] | The air inlet filter is designed such that an unclean state can be easily detected |  | 2 | 1 | 1 | 1
|  |  | Use Error - Dirty filter installed backwards |  | N | 1930 | Symptomatic; IV antibiotics required with increased ventilatory support. An infection that requires more advanced medical intervention | 3 | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator. | 2 | 1 | 1 | 1
|  |  | Dual or carrying bag dirty | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD230 | [6 empty] | Description of patient circuit | 2 | 1 | 1 | 1
|  |  | inspiratory and expiratory filters dirty | Pneumonia | N | 2011 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated | 1 | PRD427 | [5 empty] | The exterior of the ventilator is cleanable with a cloth or sponge lightly moistened with a bactericide or germicide solution | WARNING: Use all cleaning solutions and products with caution. Read and follow the instructions associated with the cleaning solutions you use to clean your ventilator. Use only those solutions listed in Table 7-1. | 2 | 1 | 1 | 1
|  |  | PM schedule not adhered to |  | N | 2011 | Symptomatic; Requires minor increase in ventilatory support, Chest Physical Therapy, antibiotics | 2 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 1 | 1 | 1
|  |  | Dual or carrying bag not used |  | N | 2011 | Symptomatic; Requires major increased ventilatory support, Antibiotic therapy, and one or more of the following. Chest Physical Therapy and nebulizer around the clock | 3 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 1
|  |  | Dirt/debris transerred from child, pet or pests |  | N | 2011 | Requires change in cardiopulmonary support such as ECMO, HFV | 4 | PRD256 | [6 empty] | WARNING: Avoid using the ventilator, if possible, in dusty environments. Dusty environments may require more vigilant monitoring, cleaning, and/or replacement of air intake and other filters. | 3 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD302 PRD360 | [5 empty] | The transport bag (dual bag) is labeled with washing machine cleaning instructions | WARNING: Regularly clean the ventilator’s Dual Bag according to manufacturer’s recommendations.
[5 empty] | N |  |  |  | PRD231 | [6 empty] | WARNING: The patient circuit is intended for single use by a single patient and should be changed according to the manufacturer’s recommendations and according to the patient circuit lifetime. Refer to the instructions for use supplied by the manufacturer of the patient circuit (included with the ventilator) and Chapter 4, “Installation and Assembly.”
[5 empty] | N | [10 empty] | Consumables and Replacement Intervals - When used under normal circumstances – a relatively dust-free atmosphere, and without damage to the device and its components (shocks, cracks, significant dirt) – the intervals for replacing the ventilator’s consumable elements are as follows: Table 10-1
[5 empty] | N | [10 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits.
[5 empty] | N |  |  |  | PRD427 | [6 empty] | Before cleaning the ventilator, first disconnect the ventilator and the patient circuit If the ventilator is used indoors, the condition of the air inlet filter should be checked monthly. If the ventilator is used outdoors or in a dusty environment, the filter should be checked weekly and replaced as necessary Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else.
VH29 | Damage to internal system components | Internal device system component damage due to oxidation, water ingress, servicing, installation. | User Error - Physical damage from liquid ingress due to spillage or excessive cleaning liquid | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD1 |  |  |  |  | The ventilator complies with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the following product Classifications as detailed in IEC 60601-1: • IPX1 with respect to ingress of moisture 560 Only: • IPX2 with respect to ingress of moisture The following cleaning agents were assessed for cleanability (X50 cycles): 1. Mild detergent (Joy) 2. 70% isopropyl alcohol (rubbing alcohol) 3. household bleach (10% Clorox or equivalent, 90% tap water) 4. Glutaraldehyde (Cidex or equivalent) 5. Amphyl 6. Cavacide 7. Hydrogen peroxide 8. Ammonia (15%) 9. Windex 10. Control III 11. Saline 12. Formula 409 Cleaner • Max volume of available liquid that could cause spillage is defined as 200mls. | - | Manual - Index of Protection rating for the ventilator’s enclosure, defined in IEC 60529 (BSEN60529). The first digit, 3, indicates protection against the intrusion of small foreign bodies (including fingers, tools, wires, etc. with a diameter greater than 2.5 mm) into the ventilator. The second digit, 1, indicates protection against water dripping or falling vertically, as well as an environment featuring water vapour condensation and/or light rain. This rating appears on the ventilator’s back panel; see Table 1-2, item 5. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use Error - ventilator is immersed in liquid |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD24 | [6 empty] | WARNING: Use all cleaning solutions and products with caution. Read and follow the instructions associated with the cleaning solutions you use to clean your ventilator. Use only those solutions listed in Table 7-1. | 2 | 1 | 1 | 1
|  |  | Use Error - Exhalation valve applied in incorrect orientation (open end facing upwards), allowing moisture enter the ventilator |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD364 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [7 empty] | WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD255 | [6 empty] | WARNING: To minimize the risk of damage, you must use the Dual Bag to transport the Puritan Bennett™ 560 Ventilator. See Figure D-2. Dual Bag.
[5 empty] | N |  |  |  | PRD440 | [6 empty] | WARNING: The ventilator should never be immersed in any liquid, and any liquid on The surface of The device should be wiped away immediately
[5 empty] | N |  |  |  | PRD236 | [6 empty] | WARNING: Do not operate the ventilator in direct sunlight, near heat sources, outdoors, or near installations where liquid may pose a risk without first providing adequate protection for the device.
[5 empty] | N |  |  |  | PRD51 | [6 empty] | PB560 WARNING: To prevent any interference with the internal sensors of the ventilator, do not install a humidifier upstream of the ventilator.
[5 empty] | N |  |  |  | PRD440 | [6 empty] | WARNING: To avoid damage to the ventilator, in particular the batteries or electrical components, fluids must not be allowed to enter the device, particularly through the air inlet filter or the cooling apertures located in the side, rear, and bottom panels of the ventilator.
VH30 | Device/Device Accessories Become Hot | Any part of device or accessories, where the user handles the device, is hot due to environmental exposure or excessive heating of internal components (e.g. batteries, electrical components). | Excessive power dissipation | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The duration of touch to ventilator and/or accessories is defined as less than 10 seconds |  |  | 2 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | power surge |  | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PRD200 | [5 empty] | HIGH / LOW BATTERY TEMP alarm HIGH INT TEMP COOL VENT alarm |  | 2 | 1 | 1 | 1
|  |  | ventilator is positioned in direct sunlight or heat source | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD51 |  |  |  |  | Safety net |  |  | 2 | 1 | 1 | 1
|  |  | Internal battery overheats |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD438 |  |  |  |  | Monitor for internal temperature sensor |  | WARNING: In case of operation in a high ambient temperature, handle the ventilator with care; some portions of the device may have high surface temperatures. WARNING: Handle the ventilator with care during and after use, particularly when ambient temperatures are high. Some ventilator surfaces may become hot, even if safety specifications are not exceeded. WARNING: In the case of high ambient temperatures, it may take a significant period of time to cool the internal temperature of the ventilator to the proper operating range. To avoid injury to the patient, ensure that the air inspired by the patient does not exceed 41°C (106°F). If in doubt, replace the ventilator. | 2 | 1 | 1 | 1
|  |  | Overcharge of the internal battery |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [5 empty] | Design - Cooling fan | COOLING FAN alarm | WARNING: Do not operate the ventilator in direct sunlight, near heat sources, outdoors, or near installations where liquid may pose a risk without first providing adequate protection for the device. Installing the ventilator - Choose an area where air can circulate freely. Avoid proximity to loose fabrics, such as curtains, and direct exposure to sunlight. | 2 | 1 | 1 | 1
|  |  | Internal battery inserted with polarity reversed |  | N |  |  |  | PRD442 | [5 empty] | TURB OVERHEAT alarm | WARNING: If the ambient temperature where the device is operated is greater than 35 °C (95 °F), the temperature of the patient circuit or the flow supplied at the device outlet may exceed 41 °C (106 °F), and the patient circuit may reach up to 60°C (140°F). This may lead to undesirable side effects for the patient. To avoid injury to the patient move the patient and the ventilator to a cooler location. For more information, contact Covidien. WARNING: Operating speed of the cooling fan not properly adjusted for the internal ambient temperature of the device. Restart ventilator to see if alarm clears. If not, replace the ventilator and call your customer service representative.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD438 | [6 empty] | WARNING: Handle the ventilator with care after use, particularly when ambient temperatures are high. Some ventilator surfaces may be very hot, even if safety specifications are not exceeded.
[5 empty] | N |  |  |  | PRD438 | [6 empty] | WARNING: In the case of high ambient temperatures, it may take a significant period of time to cool the internal temperature of the ventilator to the proper operating range. To avoid injury to the patient, ensure that the air inspired by the patient does not exceed 41°C (106°F). If in doubt, replace the ventilator.
[5 empty] | N | [8 empty] | Design- the battery has a thermal cut off and temperature sensing. The battery will turn off if it overheats.
[5 empty] | N | [8 empty] | Circulation holes in the powerpack housing to allow air to circulate
[5 empty] | N | [8 empty] | Design - Internal batteries have over voltage and over current to prevent overcharge. Under voltage and under current to prevent incorrect discharge
[5 empty] | N | [8 empty] | Design - internal battery will only make connection in one orientation. | Protection - alarm if internal battery not detected.
VH31 | Display / GUI - Blank | Display is blank during use or during service mode. | DFMEA CPU PCBA - Display Cables - Display and backlight cable | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | [6 empty] | Protective - Audio alarms will continue to annunciate |  | 1 | 2 | 1 | 1 | P2 scores are based on Clinical expertise, engineering inputs and complaint data. Ventilator transfer is likely result as ventilation continues. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD50 | [5 empty] | POST |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD417 | [5 empty] | operational verification checklist |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD32 | [5 empty] | Protective - alarm volume is adjustable 65 to 80 dBA through the preference menu |  | 1 | 3 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD296 | [6 empty] | Manual - Operational Verification Checklist - Verify the alarm volume is adapted to the patient environment. | 1 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 3 | 1 | 2
VH32 | Inability to change settings | Display/Guy or hard keys unresponsive during use or during service mode. | DFMEA CPU PCBA - Display KEYBOARD MEMBRANE - USER INTERFACE | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | [6 empty] | Audio alarms will continue to annunciate |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA UI - Incorrect Settings change |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD50 | [5 empty] | POST |  | 2 | 2 | 1 | 1
|  |  | Use error - using sharp objects (pen, sharp finger nails) to manipulate the buttons |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD32 | [5 empty] | Protective - alarm volume is adjustable 65 to 80 dBA through the preference menu |  | 2 | 1 | 1 | 1
|  |  | User misinterruption of ventilator icons |  | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD417 | [5 empty] | operational verification checklist |  | 2 | 1 | 1 | 2
|  |  |  | Anxiety | Y | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms |  | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm |  | 2 | 1 | 1 | 1
[5 empty] | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD173 | [5 empty] | HIGH RATE alarm
[5 empty] | Y |  |  |  | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms
[5 empty] | N | [9 empty] | Graphics screen or active screen remains visible.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH33 | Electrical shock | The passage of electric current through the body (e.g. exposed electrical components, transmission of electrical current through fluids including bodily). | Case opened by unqualified person | Electrical shock | N | 2554 | Asymptomatic, intervention not indicated. Slight tingling sensation; if burn see Harm of Burn | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Liquid ingress into case leading to exposure to hazardous voltages |  | N | 2554 | Non-urgent medical intervention indicated. Slight shock felt; not painful but disturbing. Average individual can let go. However, strong involuntary reactions to shocks in this range may lead to injuries .5mA ; if burn see Harm of Burn | 2 | [5 empty] | Design - battery cover requires a torque driver to remove screws Design - electric cables are insulated and no connections are able to be touched if plugged into a power source housing is non conductive Design - housing is non conductive |  |  | 1 | 1 | 1 | 1
|  |  | AC/DC power cord is damaged |  | N | 2554 | Medical intervention indicated (planned cardioversion)Painful shock, begin to lose muscular control. Commonly referred to as the freezing current or "let-go" range. 6mA to 16mA.; if burn see Harm of Burn | 3 | PRD257 PRD268 | [6 empty] | WARNING: Do not attempt to open, repair or otherwise service the ventilator yourself. Doing so might endanger the patient, damage the ventilator, and/or void your warranty. Only personnel authorized and qualified by Covidien should repair, open or service the ventilator. | 1 | 1 | 1 | 1
[5 empty] | N | 2554 | Life-threatening consequences; hemodynamic compromise; urgent intervention indicated; Extreme pain, respiratory arrest, severe muscular contractions. Individual cannot let go. Death is possible. 17mA-99mA; if burn see Harm of Burn | 4 | PRD61 | [5 empty] | Once in “Service State” the only way to exit is to power down the device. No switch from “Service State” to other states (Active ventilation State or Ventilation Stand-By State), or from other states to “Service State”, is allowed. | Service Manual - Warning - To prevent electrical shock hazard and possible personal injury, always disconnect electrical power sources before servicing the ventilator. Follow accepted safety practices for electrical equipment when testing or making equipment, adjustment, or repairs. To avoid an electrical shock hazard while servicing the ventilator, be sure to remove all power to the ventilator by disconnecting the power source and turning off all ventilator power switches before servicing. Follow accepted safety practices for electrical equipment when testing or making equipment adjustment, or repairs. | 1 | 1 | 1 | 2
[5 empty] | N | 2554 | Death ->2000mA | 5 | PRD241 PRD242 PRD243 | [6 empty] | WARNING: The power supply to which the ventilator is connected (both AC and DC) must comply with all applicable standards and provide electrical power corresponding to the voltage characteristics inscribed on the rear of the ventilator to ensure correct operation. Refer also to the electrical specifications found in Appendix B, “Specifications”. | 1 | 1 | 1 | 2
|  |  |  | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PRD258 | [6 empty] | WARNING: Ensure that the AC power cable is in perfect condition and not compressed. The device should not be turned on if the AC power cable is damaged. | 1 | 2 | 1 | 1
[5 empty] | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PRD269 | [5 empty] | An Internal label is applied to warn about high voltage. |  | 1 | 1 | 1 | 1
[5 empty] | N | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | PRD424 (560) | [6 empty] | WARNING: To connect the ventilator to an external power source, first ensure the ventilator’s I/O switch is off (O). Then, connect the desired power cable to the ventilator. Finally, connect the power cable to the external power source. WARNING: Connect the external DC power source by first connecting the power cable to the ventilator and then to the external DC source. Follow the reverse procedure to disconnect the device from the external DC power source. WARNING: To disconnect the ventilator from an external power source, first power-down the ventilator. Then, disconnect the power cable from the external power source and, finally, the ventilator. | 1 | 1 | 1 | 1
[5 empty] | N |  |  |  | PRD440 |  |  |  |  | The ventilator user and clinician manuals shall include a warning instructing the user to never immerse the ventilator in liquid or allow any liquid to enter any device opening.
[5 empty] | N |  |  |  | PRD445 |  |  |  |  | The ventilator user and clinician manuals shall include a warning instructing the user to never expose the battery or any electrical part to water.
[5 empty] | N |  |  |  | PRD226 |  |  |  |  | Design - The ventilator provides an external polarized AC cable to operate the ventilator from AC mains power.
[5 empty] | N |  |  |  | PRD428 | [6 empty] | WARNING: Do not leave power cables lying on the ground where they may pose a hazard.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH34 | Environmental specifications exceeded | Ventilator used outside of the environmental specifications as described in the user manual | User Error - environmental temperature or relative humidity exceeded | Hyperventilation | N | 1910 | Asymptomatic; PaCO2 or ETCO2 low. Mild tachypnea. | 1 | PRD130 |  |  |  |  | Design - Altitude compensation |  |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Altitude change. (excessive ventilation or not sufficient ventilation) |  | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD406 | [6 empty] | WARNING: The default setting for altitude compensation is YES. Altitude compensation should always be set to YES for accurate volume delivery calculations at all elevations. | 1 | 1 | 1 | 1
[5 empty] | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD14 | [6 empty] | Environmental The following environmental conditions shall be observed: Table A-11. Environmental Conditions for Storage or Transport Table A-12. Environmental Conditions for Operation Under extreme conditions of use, within the limits of a supply voltage of –20% and temperatures ranging from normal to 45 °C (113 °F) with ≤ 75% RH, the ventilator should not malfunction nor endanger the user. However, operating the device for prolonged periods or repeatedly under such extreme conditions could result in premature aging of components and more frequent maintenance. | 1 | 1 | 1 | 1
|  |  |  | Lung, Over inflation | N | 2397 | Asymptomatic-Observed | 1 | PRD337 | [6 empty] | Manuals contain environmental specifications for transport and storage | 1 | 1 | 1 | 1
[5 empty] | N | 2397 | Small ↑ mean airway pressure, requiring minor change in ventilator settings. | 2 | PRD238 | [6 empty] | WARNING: To ensure correct and lasting operation of the device, ensure that the ventilator is installed and operated in the environmental conditions recommended in Appendix A, “Specifications.” | 1 | 1 | 1 | 1
[5 empty] | N | 2397 | Moderate ↑ mean airway pressure, radiographical signs, requiring major change in ventilator settings. | 3 | PRD441 | [6 empty] | WARNING: If the ventilator has been transported or stored at a temperature that differs more than 20 °C ( 36 °F) from the temperature in which it will be operating, the ventilator should be allowed to stabilize in its operating environment for at least two (2) hours prior to use. | 1 | 1 | 1 | 1
[5 empty] | N | 2397 | Critical↑ mean airway pressure, radiographical signs parenchemal destruction (e.g. blebs), requiring major change in ventilator therapy (e.g., ECMO, HFV) | 4 | PRD236 | [6 empty] | Manual - WARNING: Do not operate the ventilator in direct sunlight, near heat sources, outdoors, or near installations where liquid may pose a risk without first providing adequate protection for the device. | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 1 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 1 | 1 | 1
VH35 | High PEEP | PEEP is higher than prescribed setting | User Error - Incorrect settings for prescribed ventilation | Lung, Over inflation | N | 2397 | Asymptomatic-Observed | 1 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA Breath Delivery Software target pressures and flows are too aggressive for patient Software did not detect an exhalation valve present when there is a physical exhalation valve installed Exhalation Valve control failure Alarms No detection of "Check Proximal Line 1 " alarm UI - Incorrect Settings change |  | N | 2397 | Small ↑ mean airway pressure, requiring minor change in ventilator settings. | 2 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 2 | 2 | 1 | 1
|  |  | Exhalation Bacteria filter |  | N | 2397 | Moderate ↑ mean airway pressure, radiographical signs, requiring major change in ventilator settings. | 3 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 1 | 1 | 1
[5 empty] | Y | 2397 | Critical↑ mean airway pressure, radiographical signs parenchemal destruction (e.g. blebs), requiring major change in ventilator therapy (e.g., ECMO, HFV) | 4 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 2 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD128 |  |  |  |  | PEEP pressure control accuracy during exhalation steady phase |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD74 |  |  |  |  | PEEP adjustment range |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD102 |  |  |  |  | APNEA TIME adjustment range |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD388 |  |  |  |  | Inspiratory Sensitivity settings |  |  | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD173 | [5 empty] | HIGH RATE alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD127 | [6 empty] | PB560 - WARNING: Adding attachments to the ventilator breathing system can cause the pressure during exhalation at the patient connection port to increase.
VH36 | Expiratory time too long | Expiratory time does not meet patients need | User Error - Incorrect settings for prescribed ventilation | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD146 |  |  |  |  | IE ratio |  |  | 2 | 2 | 1 | 1 | P2 scores are Based on clinical opinion and engineering support | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD388 |  |  |  |  | Inspiratory Sensitivity settings |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD87 |  |  |  |  | Exhalation sensitivity settings |  |  | 2 | 1 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD76 |  |  |  |  | Adjustable breath rate |  |  | 2 | 1 | 1 | 2
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD173 | [5 empty] | HIGH RATE alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 1 | 1 | 1
[5 empty] | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | PRD433a | [6 empty] | WARNING: The ventilator offers a variety of breath delivery options. Throughout the patient’s treatment, the clinician should carefully select the ventilation mode and settings to use for that patient, based on clinical judgment, the condition and needs of the patient, and the benefits, limitations, and characteristics of the breath delivery options. As the patient’s condition changes over time, periodically assess the chosen modes and settings to determine whether those are best for the patient’s current needs. | 2 | 1 | 1 | 2
|  |  |  | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | 1916 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) |  | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 2 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH37 | Deleted
VH38 | Fire Hazard | Fire associated with the device or accessory | DFMEA Power Supply PCBA - ACDC supply O2 SUPPLY - O2 valve oxygen connector - Connecting to the ventilator | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PRD 1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. • Not suitable for use in the presence of flammable anesthetic mixtures | - | Manual - WARNING: The ventilator must not be used with flammable anesthetic substances | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | battery |  | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PRD444 | [6 empty] | WARNING: Never expose any batteries to direct flame. | 1 | 2 | 1 | 1
|  |  | electrical sparking with an oxygen rich environment |  | N | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | [7 empty] | WARNING: To reduce the risk of a fire hazard, keep matches, lighted cigarettes, and all other sources of ignition (such as flammable anesthetics and/or heaters) away from the ventilator and oxygen hoses. | 1 | 1 | 1 | 1
[5 empty] | N | 2530 | Thermal injury not recognized at the time of surgery. Results in the patient having to undergo urgent surgical intervention due to leakage of a structure/organ. (May need to see peritonitis, aroma, sepsis, etc.) | 4 | PRD439 | [6 empty] | WARNING: Even though the Puritan Bennett™ 560 Ventilator meets current safety standards, the internal Lithium-ion battery of the device exceeds the 100Wh threshold and is therefore considered to be Dangerous Goods (DG) Class 9 – Miscellaneous, when transported in commerce. As such, the Puritan Bennett™ 560 Ventilator and/or the associated Lithium-ion battery are subject to strict transport conditions under the Dangerous Goods Regulation for air transport (IATA: International Air Transport Association), International Maritime Dangerous Goods code for sea and the European Agreement concerning the International Carriage of Dangerous Goods by Road (ADR) for Europe. Private individuals who transport the device are excluded from these regulations although for air transport some requirements apply. For air transport; the Puritan Bennett™ 560 Ventilator is permitted as checked-in or carry-on baggage. Two spare batteries per person may be taken on board as carry-on luggage only, with the prior approval of the airline. This classification and regulatory requirements may vary depending upon the country and mode of transport. Therefore it is recommended that users verify with the carrier / airline as to which measures to take before the voyage. | 1 | 1 | 1 | 2
[5 empty] | N | 2530 | Death | 5 | PRD253 | [6 empty] | WARNING: To avoid injury to the patient and/or possible damage to the ventilator: before using the ventilator, use a flow meter (flow regulator) to regulate the oxygen supply to specifications before connecting the ventilator to the oxygen supply. The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD421 | [6 empty] | WARNING: The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. Inspect the oxygen coupler (Figure 6-16, item 2) before use to ensure it has its black O-ring attached and in good condition. Do not use an oxygen coupler with a missing, damaged, or worn O-ring. Ensure that the oxygen supply pressure to the machine never exceeds 7 psi (50 kPa) or a flow of 15 lpm. Refer to Table A-8. on page A-4 for volume and sensitivity tolerances. | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD429 | [6 empty] | WARNING: The coupler must not remain connected to the oxygen connector unless it also connected to a leak-proof, external oxygen gas source. When an oxygen supply is not being used with the ventilator, disconnect the oxygen source completely from the ventilator. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD229 | [6 empty] | PB560 - WARNING: In the event of an oxygen leak, shut down the supply of oxygen at its source. In addition, remove and/or keep any incandescent source away from the device, which may be enriched with oxygen. Circulate fresh air into the room to bring the oxygen level down to normal.
[5 empty] | N |  |  |  | PRD421 | [6 empty] | PB560 - WARNING: The oxygen supply hose ages even when it is not in use and should be replaced periodically. Follow the expiration date, if any.
[5 empty] | N |  |  |  | PRD157 | [6 empty] | PB560 - WARNING: The oxygen supply must be shut off when ventilation is interrupted. Before disconnecting the oxygen hose, allow the ventilator to continue for a few cycles without oxygen to flush the patient circuit of excess oxygen.
[5 empty] | N |  |  |  | PRD421 | [6 empty] | PB560 - WARNING: Before connecting the oxygen supply, ensure that the stud on the oxygen connector is protruding outwards.
[5 empty] | N |  |  |  | PRD303 | [6 empty] | WARNING: Ensure device is not positioned or located such that the AC and DC connections at the back of the vent are difficult to access.
VH39 | Hazardous materials | Device or device materials when discarded may cause effects to patient, user or environment. | Improper battery disposal | Chemical exposure, Hazardous | N | 2570 | Minor Inflammation, no treatment needed | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. |  |  | 3 | 1 | 1 | 1 | P2 scores are based on clinical opinion, engineering inputs and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Improper device disposal |  | N | 2570 | Temporary injury or impairment that requires minor professional medical intervention | 2 | PRD374 |  |  |  |  | The ventilator complies with directive 2006/66/EC on batteries and accumulators and waste batteries and accumulators, 2006. |  | WARNING: For environmental protection, the ventilator and its components, whatever their respective conditions of operation, cannot be disposed of with household waste and must be submitted for suitable selective collection and possible recycling. Observe all applicable regulations when disposing of the ventilator and any of its components. | 3 | 1 | 1 | 1
|  |  | Choice of materials at design stage |  | N | 2570 | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | PRD375 |  |  |  |  | The ventilator complies with directive 2002/96/EC, 2003 on waste electrical and electronic equipment (WEEE), 2003. |  | Warnings Regarding the Environment • For environmental protection, the ventilator and its components, whatever their respective conditions of operation, cannot be disposed of with household waste and must be submitted for suitable selective collection and possible recycling. Observe all applicable regulations when disposing of the ventilator and any of its components. • Follow local governing ordinances and recycling regulations regarding disposal or recycling of batteries and other device components. | 3 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD373 PRD265 PRD266 PRD270 | [5 empty] | The ventilator and battery are labeled to indicate that the device shall be considered as waste electrical and electronic equipment. The manuals indicate that the devices should be considered as waste electrical and electronic equipment. | Recycle electrical and electronic items with this marking according to local regulations (EN50419). Do not discard with ordinary household waste. This includes the ventilator and the battery. The battery is labeled to indicate that the battery shall be considered as waste electrical and electronic equipment. | 3 | 1 | 1 | 1
VH40 | Humidifier failure - dry | External humidifier delivers humidification that is dry e.g. no moisture delivered to patient | Use Error - Humidifier runs dry | Pneumothorax | N | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated Asymptomatic-clinical or diagnostic finding. | 1 | PRD412 | [6 empty] | WARNING: Always position a humidification device so that it is lower than both the ventilator and the patient. Use water traps, if necessary, to limit water in the patient circuit and precautions should be taken when discarding the fluid in the water trap, discard per local ordinance for proper disposal. If a heated humidifier is used, you should always monitor the temperature of the gas delivered to the patient. Gas delivered from the ventilator that becomes too hot may burn the patient's airway. | 1 | 1 | 1 | 1 | P2 scores are based on clinical opinion, engineering inputs and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Heater failure |  | N | 2012 | Symptomatic intervention indicated (e.g. tube placement without sclerosis (no pleurodesis), hospitalization observation, electively place chest tube or | 2 | PRD235 | [6 empty] | WARNING: During invasive ventilation (when an artificial airway bypasses the patient’s upper respiratory system), the patient’s upper respiratory system cannot humidify the incoming gas. For this reason, the use of a humidifier, to minimise drying of the patient’s airways and subsequent irritation and discomfort, must be used. | 1 | 1 | 1 | 1
|  |  | Use Error - incorrect set up of humidifier/HME |  | N | 2012 | Sclerosis and/or operative intervention indicated; hospitalization indicated (Pleurodesis) Ventilation: Pneumothorax requiring placement chest tube with hospitalization or prolonged hospitalization | 3 | PRD413 | [6 empty] | Adding accessories to the ventilator breathing circuit, such as a humidifier and water trap(s), may result in a decrease in tidal volume delivered to the patient due to the added compressible volume of the accessory. Always assure that the patient is receiving the appropriate inspired volume when altering the breathing circuit configuration. | 1 | 1 | 1 | 1
|  |  | HME fitted at the incorrect location within the breathing circuit | Dehydration | N | 1807 | Dry skin; decreased skin turgor. Need for increased oral fluids. Ventilation: Asymptomatic dehydration, normal. Dry Mucosa. Need to add topical hydration, humidity. | 1 | PRD173 | [5 empty] | HIGH RATE alarm |  | 1 | 1 | 1 | 1
[5 empty] | Y | 1807 | IV fluids indicated < 24 hours. Ventilation: Drying of airway requiring lavage and suction. | 2 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 1 | 1 | 1 | 1
[5 empty] | N | 1807 | IV Fluids indicated > 24 hours; hospitalization required. Ventilation: Drying of airway requiring lavage, mucolytics, and suction. | 3 | [6 empty] | Humidifiers have temperature and functional alarms. |  | 1 | 1 | 1 | 1
[5 empty] | N | 1807 | Life threatening consequences; patient requiring support for heart rate, blood pressure, kidney function including dialysis (May also need to see hypovolemic shock). Hardened, partial and /or Complete airway obstruction requiring bronchoscopy. | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 1 | 1 | 1
|  |  |  | Hyperventilation |  | 1910 | Asymptomatic; PaCO2 or ETCO2 low. Mild tachypnea. | 1 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires minor ventilator adjustment | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires major ventilator adjustment | 3 | [8 empty] | 1 | 1 | 1 | 1
VH41 | High inspiratory gas pressure | Pressure delivered to the patient is too high | User Error - Incorrect settings for prescribed ventilation | Lung, Over inflation | Y | 1916 | Asymptomatic-Observed | 1 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, complaint data, and assumed underreporting. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA - Barometric pressure sensor '+5V Reference Internal pressure sensor measure Inspiratory flow measure |  | N | 1916 | Small ↑ mean airway pressure, requiring minor change in ventilator settings. | 2 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 1 | 2 | 1 | 1
|  |  | SWFMEA Breath delivery Software target pressures and flows are too aggressive for patient Software did not detect an exhalation valve present when there is a physical exhalation valve installed Exhalation Valve control failure Alarms - No detection of High Peak Pressure No detection of "Check Proximal Line 1 " alarm UI - incorrect settings change |  | N | 1916 | Moderate ↑ mean airway pressure, radiographical signs, requiring major change in ventilator settings. | 3 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 1 | 1 | 1 | 1
|  |  | Circuit obstruction/occlusion | Pneumothorax | N | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated Asymptomatic-clinical or diagnostic finding. | 1 | PRD162 | [5 empty] | Graphics to display patient response - |  | 1 | 2 | 1 | 1
|  |  | Excessive patient secretions |  | N | 2012 | Small to Moderate Pneumothorax, electively place chest tube or observation | 2 | PRD146 |  |  |  |  | The ventilator displays the ratio of inspiratory time to expiratory time |  |  | 1 | 2 | 1 | 1
|  |  | Exhalation valve malfunction |  | N | 2012 | Large Pneumothorax, electively place chest tube | 3 | PRD133 | [5 empty] | Device truncates breath and transitions to exhalation limiting time at high pressure. |  | 1 | 1 | 1 | 1
|  |  | Use Error - incorrect set up of humidifier |  | Y | 2012 | Tension Pneumothorax; Emergent Chest tube placement, intubation. | 4 | PRD179 (560) | [5 empty] | LOW VTE Alarm |  | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD128 |  |  |  |  | PEEP pressure control accuracy during exhalation steady phase
[5 empty] | N |  |  |  | PRD173 | [5 empty] | HIGH RATE alarm
[5 empty] | Y |  |  |  | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms
[5 empty] | Y |  |  |  | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms
[5 empty] | Y |  |  |  | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm
[5 empty] | N |  |  |  | PRD74 |  |  |  |  | PEEP adjustment range
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD138 | [6 empty] | WARNING: Use of a nebulizer or humidifier can lead to an increase in the resistance of inspiratory and exhalation filters. Monitor the filters frequently for increased resistance or blockage
VH42 | High inspiratory flow | Flow delivered to the patient is higher than set or expected | User Error - Incorrect settings for prescribed ventilation | Hyperventilation | Y | 1910 | Asymptomatic; PaCO2 or ETCO2 low. Mild tachypnea. | 1 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 1 | 5 | 3 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, literature, Clinical evaluation reports (CER). assumed underreporting | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA - Barometric pressure sensor Inspiratory flow measure Inspiratory flow sensor fitting |  | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD177 |  |  |  |  | The ventilator resets the high inspiratory pressure alarm condition on the first breath that does not exceed the alarm threshold, but the visual alarm persists until the operator presses the Alarm control key. |  |  | 1 | 2 | 1 | 1
|  |  | SWFMEA Breath Delivery Software target pressures and flows are too aggressive for patient Software did not detect an exhalation valve present when there is a physical exhalation valve installed Exhalation Valve control failure |  | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 1 | 1 | 1 | 1
|  |  | Alarms No detection of High Peak Pressure |  | N | 1910 | Symptomatic; PaCO2 or ETCO2 unacceptable range, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD162 | [5 empty] | Graphics to display patient response - |  | 1 | 1 | 1 | 2
|  |  | UI - Incorrect Settings change |  | N | 1910 | Death | 5 | PRD146 |  |  |  |  | The ventilator displays the ratio of inspiratory time to expiratory time |  |  | 1 | 1 | 1 | 2
|  |  |  | Lung, Over inflation | N | 2397 | Asymptomatic-Observed | 1 | PRD133 | [5 empty] | Device truncates breath and transitions to exhalation limiting time at high pressure. |  | 1 | 2 | 1 | 1
[5 empty] | Y | 2397 | Small ↑ mean airway pressure, requiring minor change in ventilator settings. | 2 | PRD179 (560) | [5 empty] | LOW VTE Alarm |  | 1 | 2 | 1 | 1
[5 empty] | N | 2397 | Moderate ↑ mean airway pressure, radiographical signs, requiring major change in ventilator settings. | 3 | PRD173 | [5 empty] | HIGH RATE alarm |  | 1 | 1 | 1 | 1
[5 empty] | N | 2397 | Critical↑ mean airway pressure, radiographical signs parenchemal destruction (e.g. blebs), requiring major change in ventilator therapy (e.g., ECMO, HFV) | 4 | PRD128 |  |  |  |  | PEEP pressure control accuracy during exhalation steady phase |  |  | 1 | 1 | 1 | 2
[5 empty] | Y | 2397 | Death | 5 | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms |  | 1 | 1 | 1 | 2
[5 empty] | Y |  |  |  | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms
[5 empty] | N |  |  |  | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold
[5 empty] | Y |  |  |  | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm
[5 empty] | N |  |  |  | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH43 | High Inspiratory gas temperature | High inlet gas to the vent causing rise to inspired gas temperature | DFMEA humidifier - heater | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | [5 empty] | Cooling fan |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, literature, Clinical evaluation reports (CER). assumed underreporting | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use environment is a high temperature |  | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PRD200 | [5 empty] | COOLING FAN alarm HIGH INT TEMP COOL VENT alarm TURB OVERHEAT alarm HIGH / LOW BATTERY TEMP alarm |  | 2 | 2 | 1 | 1
[5 empty] | N | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | PRD438 |  |  |  |  | Monitor for internal temperature sensor |  | WARNING: In case of operation in a high ambient temperature, handle the ventilator with care; some portions of the device may have high surface temperatures. WARNING: Handle the ventilator with care during and after use, particularly when ambient temperatures are high. Some ventilator surfaces may become hot, even if safety specifications are not exceeded. | 2 | 1 | 1 | 1
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD442 | [6 empty] | WARNING: If the ambient temperature where the device is operated is greater than 35 °C (95 °F), the temperature of the patient circuit or the flow supplied at the device outlet may exceed 41 °C (106 °F), and the patient circuit may reach up to 60°C (140°F). This may lead to undesirable side effects for the patient. To avoid injury to the patient move the patient and the ventilator to a cooler location. For more information, contact Covidien. | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD236 | [6 empty] | WARNING: Do not operate the ventilator in direct sunlight, near heat sources, outdoors, or near installations where liquid may pose a risk without first providing adequate protection for the device. Installing the ventilator - Choose an area where air can circulate freely. Avoid proximity to loose fabrics, such as curtains, and direct exposure to sunlight. | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD442 | [6 empty] | WARNING: If the ambient temperature where the device is operated is greater than 35 °C (95 °F), the temperature of the patient circuit or the flow supplied at the device outlet may exceed 41 °C (106 °F), and the patient circuit may reach up to 60°C (140°F). This may lead to undesirable side effects for the patient. To avoid injury to the patient move the patient and the ventilator to a cooler location. For more information, contact Covidien. | 2 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD412 | [6 empty] | WARNING: Always position a humidification device so that it is lower than both the ventilator and the patient. Use water traps, if necessary, to limit water in the patient circuit and precautions should be taken when discarding the fluid in the water trap, discard per local ordinance for proper disposal If a heated humidifier is used, you should always monitor the temperature of the gas delivered to the patient. Gas delivered from the ventilator that becomes too hot may burn the patient's airway. | 2 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH44 | High inspiratory volume | Volume delivered to the patient is higher than set | User Error - Incorrect settings for prescribed ventilation | Hyperventilation | Y | 1910 | Asymptomatic; PaCO2 or ETCO2 low. Mild tachypnea. | 1 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 2 | 3 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, literature, Clinical evaluation reports (CER). assumed underreporting, and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA - Barometric pressure sensor Exhalation flow measure |  | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD177 | [5 empty] | The ventilator resets the high inspiratory pressure alarm condition on the first breath that does not exceed the alarm threshold, but the visual alarm persists until the operator presses the Alarm control key. |  | 2 | 3 | 1 | 1
|  |  | SWFMEA Breath-Delivery Software target pressures and flows are too aggressive for patient Software did not detect an exhalation valve present when there is a physical exhalation valve installed Exhalation Valve control failure Alarms No detection of High Peak Pressure Incorrect monitored exhaled tidal volume (VTE) from flow sensor UI - Incorrect Settings change |  | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1
[5 empty] | N | 1910 | Symptomatic; PaCO2 or ETCO2 unacceptable range, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 2 | 2 | 1 | 2
[5 empty] | N | 1910 | Death | 5 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 2 | 1 | 2
|  |  |  | Lung, Over inflation | Y | 2397 | Asymptomatic-Observed | 1 | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms |  | 2 | 3 | 1 | 1
[5 empty] | N | 2397 | Small ↑ mean airway pressure, requiring minor change in ventilator settings. | 2 | PRD173 | [5 empty] | HIGH RATE alarm |  | 2 | 3 | 1 | 1
[5 empty] | Y | 2397 | Moderate ↑ mean airway pressure, radiographical signs, requiring major change in ventilator settings. | 3 | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2397 | Critical↑ mean airway pressure, radiographical signs parenchemal destruction (e.g. blebs), requiring major change in ventilator therapy (e.g., ECMO, HFV) | 4 | PRD178 (560) | [5 empty] | HIGH VTE ALARM |  | 2 | 2 | 1 | 2
[5 empty] | Y | 2397 | Death | 5 | PRD 180 | [5 empty] | HIGH VTI ALARM |  | 2 | 2 | 1 | 2
[5 empty] | N |  |  |  | PRD128 |  |  |  |  | PEEP pressure control accuracy during exhalation steady phase
[5 empty] | Y |  |  |  | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms
[5 empty] | N |  |  |  | PRD74 |  |  |  |  | PEEP adjustment range
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH45 | High O2 delivery | Device delivers O2 above the prescribed oxygen setting | User Error - Incorrect settings for prescribed ventilation | Oxygen saturation, High | N | 2478 | Asymptomatic; no treatment necessary; found on arterial blood gas measurement; may require decreasing supplemental oxygen | 1 | PRD254 | [6 empty] | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FiO2 kit) that features a minimum and Installation and Assembly maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 1 | 3 | 1 | 1 | Rationale for harm and occurrence in this patient population is based on clinical expert opinion, engineering inputs, literature, Clinical evaluation reports (CER). assumed underreporting of use error for alarm settings/patient settings. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA: FIO2 monitoring Valve O2 connector O2 valve command |  | N | 2478 | Asymptomatic; Oxygen level is decreased (e.g., for retinal protection-neonate, protocol adherence) | 2 | PRD229 | [6 empty] | Description of Oxygen function and installation | 1 | 3 | 1 | 1
|  |  | DFMEA O2 SUPPLY - O2 valve Pneumatic fittings - O2 sensor |  | Y | 2478 | Symptomatic; Requires elective increase in non-invasive ventilatory support due to decrease of respiratory drive (e.g., COPD) | 3 | PRD321 (560) PRD322 (560) | [5 empty] | PB560 CHECK FIO2 SENSOR alarm LOW FIO2 alarm HIGH FIO2 alarm FIO2 SENSOR MISSING alarm CALIBRATE FIO2 alarm | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FIO2 kit) that features a minimum and maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 1 | 2 | 1 | 1
|  |  | SWFMEA No detection of FIO2 alarm condition |  | N | 2478 | Symptomatic; Requires emergent intubation or rescue due to loss of respiratory drive (e.g., COPD, congenital cardiac cyanotic lesions requiring <85% SpO2s) | 4 | PRD253 | [6 empty] | WARNING: To avoid injury to the patient and/or possible damage to the ventilator: before using the ventilator, use a flow meter (flow regulator) to regulate the oxygen supply to specifications before connecting the ventilator to the oxygen supply. The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. | 1 | 2 | 1 | 2
[5 empty] | N | 2478 | Death | 5 | PRD38 |  |  |  |  | The ventilator provides O2 intake port and connector which allows connecting of low pressure oxygen sources at less than 15 lpm at 50 KPa (e.g. Oxygen concentrator). |  |  | 1 | 2 | 1 | 2
|  |  |  | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | [7 empty] | WARNING: Oxygen therapy for patients with respiratory failure is a common and effective medical prescription. However, be aware that inappropriate oxygen use may potentially lead to serious complications, including, but not limited to, patient injury. | 1 | 3 | 1 | 1
[5 empty] | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | PRD421 | [6 empty] | WARNING: The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. Inspect the oxygen coupler (Figure 6-16, item 2) before use to ensure it has its black O-ring attached and in good condition. Do not use an oxygen coupler with a missing, damaged, or worn O-ring. Ensure that the oxygen supply pressure to the machine never exceeds 7 psi (50 kPa) or a flow of 15 lpm. Refer to Table A-8. on page A-4 for volume and sensitivity tolerances. | 1 | 3 | 1 | 1
[5 empty] | N | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Examples: Bleeding vessels requiring suture repair with blood transfusions; additional/unintended tissue resection required to achieve intended patient outcome. Procedure changed to open procedure to manage bleeding. Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | [7 empty] | WARNING: Ensure that the oxygen supply pressure to the machine never exceeds 7 psi (50 kPa) or a flow of 15 lpm. Refer to Table A-8. on page A-4 for volume and sensitivity tolerances. | 1 | 2 | 1 | 1
[5 empty] | N | 2104 | Sudden significant injury to tissue that requires immediate attention to prevent a catastrophic event; Examples: Profuse bleeding on a large vessel (pulmonary artery/vein, renal artery/vein, SMA/IMA); additional unintended tissue loss/resection needed causing additional medical intervention such as a permanent colostomy. Ventilation: Symptomatic; Signs of ARDS requiring Major change in therapy (e.g., ECMO, HFV). | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 3 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD264 |  |  |  |  | PM schedule - O2 valve |  |  | 1 | 3 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD162 | [5 empty] | graphics to display patient response |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 2 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 2 | 1 | 2
VH46 | Humidifier failure - wet | External humidifier delivers humidification that is wet e.g. excessive moisture delivered to patient | Use Error - humidifier temperature set too high/low | Obstruction | N | 2422 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated. | 1 | PRD349 | [6 empty] | WARNING: To ensure proper performance of the ventilator, use only accessories (including oxygen accessories) approved and recommended by Covidien. See Appendix H, “Parts and Accessories” or contact your customer services. | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, varied environments for use, assumed underreporting | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | humidifier temperature sensor incorrect leading to excessive moisture |  | N |  | Symptomatic, altered function, limiting of activities of daily living. | 2 | PRD412 | [6 empty] | WARNING: Always position a humidification device so that it is lower than both the ventilator and the patient. Use water traps, if necessary, to limit water in the patient circuit and precautions should be taken when discarding the fluid in the water trap, discard per local ordinance for proper disposal If a heated humidifier is used, you should always monitor the temperature of the gas delivered to the patient. Gas delivered from the ventilator that becomes too hot may burn the patient's airway. | 1 | 2 | 1 | 1
|  |  | user environment is too cold - temperature sensor on the humidifier over compensates |  | N |  | Hospitalization indicated; non-urgent elective operative intervention indicated; disabling. | 3 | PRD235 | [6 empty] | WARNING: During invasive ventilation (when an artificial airway bypasses the patient’s upper respiratory system), the patient’s upper respiratory system cannot humidify the incoming gas. For this reason, the use of a humidifier, to minimize drying of the patient’s airways and subsequent irritation and discomfort, must be used. | 1 | 2 | 1 | 1
|  |  | Use of a humidifier in combination with a HME | Hypoventilation | Y | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD190 | [5 empty] | Occlusion alarms |  | 1 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD413 | [6 empty] | WARNING: Adding accessories to the ventilator breathing circuit, such as a humidifier and water trap(s), may result in a decrease in tidal volume delivered to the patient due to the added compressible volume of the accessory. Always assure that the patient is receiving the appropriate inspired volume when altering the breathing circuit configuration. | 1 | 2 | 1 | 1
[5 empty] | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 1 | 2 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N | [9 empty] | Humidifiers have temperature and functional alarms.
[5 empty] | N |  |  |  | PRD173 | [5 empty] | HIGH RATE alarm
[5 empty] | Y |  |  |  | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm
VH47 | Improper Ergonomic Approach | Service tech or user does not use proper technique in lifting heavy ventilator component. |  | N/A this homecare ventilator is less than 5kg |  |  |  |  | PRD18 |  |  |  |  | The base model ventilator weight shall be less than 5 kg.
VH48 | Deleted
VH49 | Deleted
VH50 | Inadequate surge protection | Electrical circuit fails during use due to a power surge. | Use error - non polarized power cord used | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical experience, engineering inputs, complaint data and environment of use. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Damaged power cord used |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD241 PRD242 PRD243 | [6 empty] | WARNING: The power supply to which the ventilator is connected (both AC and DC) must comply with all applicable standards and provide electrical power corresponding to the voltage characteristics inscribed on the rear of the ventilator to ensure correct operation. Refer also to the electrical specifications found in Appendix A, “Specifications.” | 2 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD424 (560) | [6 empty] | WARNING: To connect the ventilator to an external power source, first ensure the ventilator’s I/O switch is off (O). Then, connect the desired power cable to the ventilator. Finally, connect the power cable to the external power source. WARNING: Connect the external DC power source by first connecting the power cable to the ventilator and then to the external DC source. Follow the reverse procedure to disconnect the device from the external DC power source. WARNING: To disconnect the ventilator from an external power source, first power-down the ventilator. Then, disconnect the power cable from the external power source and, finally, the ventilator. | 2 | 2 | 1 | 1
[5 empty] | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 1 | 1 | 2
[5 empty] | N | 1916 | Death | 5 | PRD226 |  |  |  |  | Design - The ventilator provides an external polarized AC cable to operate the ventilator from AC mains power. |  |  | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD227 |  |  |  |  | Design - Ventilator will transition to battery power, ventilation continues |  |  | 2 | 1 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) |  | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | 2477 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH51 | Insect Infestation | Multiple quantities of Insects found inside the Ventilator | inappropriate storage, inadequate packaging | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the following product Classifications as detailed in IEC 60601-1: • IPX1 with respect to ingress of moisture 560 Only: • IPX2 with respect to ingress of moisture | - | The ventilator will be constructed to comply with the following product Classifications as detailed in 60601-1: Protection index of enclosure: IP32 Clinician's Manual - Index of Protection rating for the ventilator’s enclosure, defined in IEC 60529 (BSEN60529:1991). The first digit, 3, indicates protection against the intrusion of small foreign bodies (including fingers, tools, wires, etc. with a diameter greater than 2.5 mm) into the ventilator. The second digit, 2, indicates protection against water dripping or falling vertically, as well as an environment featuring water vapour condensation and/or light rain. This rating appears on the ventilator’s back panel; see Table 1-2, item 5. | 1 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, literature, Clinical evaluation reports (CER). | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  |  | Infection | N | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 1 | 1 | 1
[5 empty] | N | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth or inhalation, Incision/drainage see abscess | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 1 | 1 | 1
[6 empty] | 1930 | Symptomatic; IV antibiotics required with increased ventilatory support. An infection that requires more advanced medical intervention | 3 | [8 empty] | 1 | 1 | 1 | 1
|  |  |  | anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 1 | 1 | 1 | 1
VH52 | Loss of AC external power | Loss of prescribed or ordered ventilation | A/C voltage too high or too low. Spikes in voltage. (Ventilator shut down) | Respiratory failure | Y | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD197 | [5 empty] | AC POWER DISCONNECTION alarm |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical knowledge, complaint data, engineering input, transition to battery, and patient data related to ventilator transfer. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | A/C power cord accidentally unplugged or broken (Ventilator shut down) |  | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD158 |  |  |  |  | The ventilator provides a visual indication of the active ventilator power source (AC, DC or Battery) that is operational in all ventilator states and modes. |  | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients. | 2 | 2 | 1 | 1
|  |  | DFMEA Power Supply PCBA Current control for battery charge (Constant current Generator) |  | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD299 |  |  |  |  | Design - power switch cover to prevent inadvertent power off |  |  | 2 | 1 | 1 | 2
|  |  | A/C power outage. (Ventilator shut down) |  | N | 2484 | Death | 5 | PRD241 PRD242 PRD243 | [6 empty] | WARNING: The power supply to which the ventilator is connected (both AC and DC) must comply with all applicable standards and provide electrical power corresponding to the voltage characteristics inscribed on the rear of the ventilator to ensure correct operation. Refer also to the electrical specifications found in Appendix B, “Specifications”. | 2 | 1 | 1 | 2
|  |  | Misidentification of AC power cord | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD227 |  |  |  |  | Design - Ventilator will transition to battery power, ventilation continues |  |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD194 PRD195 | [5 empty] | DEVICE FAULT 3 alarm DEVICE FAULT 4 alarm POWER FAULT alarm |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD258 | [6 empty] | Ensure that the AC and DC power cables are in perfect condition and not compressed. The device should not be turned on if the AC or DC power cable is damaged. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD243 | [6 empty] | PB560 - WARNING: The operator should connect the ventilator to an AC power source whenever available, for safer operation.
[5 empty] | N |  |  |  | PRD252 | [6 empty] | PB560 - WARNING: For the AC (“mains”) power cable to be properly secured, the attachment located on the power cable must be fitted into the power cable holder incorporated in the battery access cover and located under the AC (mains) power socket. Refer to section 4.2, “Connecting to External AC Power.”
VH53 | Loss of DC external power | Loss of prescribed or ordered ventilation. | External power source discharged or car battery not available (Ventilator shut down) | Respiratory failure | Y | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD198 | [5 empty] | DC POWER DISCONNECTION alarm |  | 2 | 1 | 1 | 1 | P2 scores are based on clinical knowledge, complaint data, engineering input, internal battery transition, and patient data related to ventilator transfer. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | External power cord accidentally unplugged or damaged (Ventilator shut down) |  | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD158 |  |  |  |  | The ventilator provides a visual indication of the active ventilator power source (AC, DC or Battery) that is operational in all ventilator states and modes. |  | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients. WARNING: When using a car auxiliary adapter (cigarette lighter) ensure the car has been started prior to plugging in the ventilator’s DC adapter. | 2 | 1 | 1 | 1
|  |  | DFMEA 12V car connection cable - Power to ventilator |  | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD299 |  |  |  |  | Design - power switch cover to prevent inadvertent power off |  |  | 2 | 1 | 1 | 2
|  |  | External power connections reversed. (Ventilator shut down) |  | N | 2484 | Death | 5 | PRD223 |  |  |  |  | The device is designed to be protected in case power connections are reversed at DC input. |  |  | 2 | 1 | 1 | 2
|  |  | Incorrect DC voltage | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD241 PRD242 PRD243 | [6 empty] | WARNING: The power supply to which the ventilator is connected (both AC and DC) must comply with all applicable standards and provide electrical power corresponding to the voltage characteristics inscribed on the rear of the ventilator to ensure correct operation. Refer also to the electrical specifications found in Appendix B, “Specifications”. | 2 | 1 | 1 | 1
|  |  | Use Error -incorrect connection of DC power |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD227 |  |  |  |  | Design - Ventilator will transition to internal battery power, Ventilation continues |  |  | 2 | 1 | 1 | 1
|  |  | Misidentification of DC power cable |  | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD194 PRD195 | [5 empty] | DEVICE FAULT 3 alarm DEVICE FAULT 4 alarm POWER FAULT alarm |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD351 |  |  |  |  | DC power cable retainer |  |  | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD258 | [6 empty] | Ensure that the AC and DC power cables are in perfect condition and not compressed. The device should not be turned on if the AC or DC power cable is damaged. | 2 | 1 | 1 | 1
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD408 | [6 empty] | WARNING: When using a car auxiliary adapter (cigarette lighter) ensure the car has been started prior to plugging in the ventilator’s DC adapter.
[5 empty] | N |  |  |  | PRD242 | [6 empty] | WARNING: Connect the external DC power source by first connecting the power cable to the ventilator and then to the external DC source. Follow the reverse procedure to disconnect the device from the external DC power source.
[5 empty] | N |  |  |  | PRD1a | [6 empty] | WARNING: Do not connect the ventilator to the battery of a battery-powered wheelchair unless the connection is listed in the instructions for use of the ventilator or the wheelchair, as this can affect the ventilator performance, which can consequently result in patient death
VH54 | Loose Internal mechanical interface or connection | Any internal parts, connectors, or components that become loose or disconnected | User Error - Physical damage from impact or vibration, or transport bag fault | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the relevant clauses of IEC 60601-1: • Class II Equipment • Internally Powered Equipment • Detachable power supply cord |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical and engineering expertise. It was considered that hazard can happen unexpectedly with unknown impact. The occurrence of harm is low for this hazard per complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Service Error |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD8 |  |  |  |  | The PB560 shall comply with Medical electrical equipment Part 2-72: Particular requirements for basic safety and essential performance of home healthcare environment ventilators for ventilator-dependent patients - EN ISO 80601-2-72. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD255 | [6 empty] | WARNING: To minimize the risk of damage, you must use the Dual Bag to transport the Puritan Bennett™ 560 Ventilator. See Figure D-2. Dual Bag. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD314 |  |  |  |  | The ventilator shall comply with sections 21.6a (vibration sinusoidal), 21.6b (random vibration) and 21.6c (bump test) of Lung ventilator for medical use - Part 3: Particular requirements for emergency transport ventilators ISO 10651-3. |  |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD364 |  |  |  |  | The ventilator shall be compatible with drop test according to 10.1.3 d article of IEC 60601-1-11 MEDICAL ELECTRICAL EQUIPMENT – PART 1-11: General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment (document IEC/SC62A 62A/624/CD) 2008. |  |  | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [6 empty] | Protect: Depending on the failure, visual and audible alarms appropriate for the condition will alarm. |  | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD257 | [6 empty] | WARNING: To ensure proper servicing and avoid the possibility of physical injury to personnel or damage to the ventilator, only personnel authorized and qualified by Covidien should attempt to service or make authorized modifications to the Puritan Bennett™ 560 Ventilator. Do not attempt to open, repair or otherwise service the ventilator yourself. Doing so might endanger the patient, damage the ventilator, and/or void your warranty. Only personnel authorized and qualified by Covidien should repair, open or service the ventilator.
[5 empty] | N | [10 empty] | WARNING: Connect the external DC power source by first connecting the power cable to the ventilator and then to the external DC source. Follow the reverse procedure to disconnect the device from the external DC power source.
[5 empty] | N |  |  |  | PRD426 | [6 empty] | WARNING: Never use a ventilator or any components or accessories that appear to be damaged. If any signs
VH55 | Loss of external mechanical interface or connection | Any external parts, connectors, or components that become disconnected | User Error - cables not connected correctly | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the relevant clauses of following product Classifications as detailed in Clause 5 of IEC 60601-1: • Class II Equipment • Internally Powered Equipment • Detachable power supply cord |  | of damage are evident, contact your equipment supplier or Covidien. | 2 | 2 | 1 | 1 | P2 scores are based on the chance that the hazard can happen anytime, anywhere and could be an unknown impact. The occurrence of harm is low for this hazard per complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User Error - Physical damage from impact or vibration, or transport bag fault |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD8 |  |  |  |  | The PB560 shall comply with Lung Ventilators for Medical Use- Particular Requirements for Basic Safety and Essential Performance Part 2: Home Care Ventilators for Ventilator-Dependent Patient EN 10651-2. |  |  | 2 | 2 | 1 | 1
|  |  | Service Error |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD255 | [6 empty] | WARNING: To minimize the risk of damage, you must use the Dual Bag to transport the Puritan Bennett™ 560 Ventilator. See Figure D-2. Dual Bag. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD314 |  |  |  |  | The ventilator shall comply with sections 21.6a (vibration sinusoidal), 21.6b (random vibration) and 21.6c (bump test) of Lung ventilator for medical use - Part 3: Particular requirements for emergency transport ventilators ISO 10651-3 |  |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD364 |  |  |  |  | The ventilator shall be compatible with drop test according to 10.1.3 d article of IEC 60601-1-11 MEDICAL ELECTRICAL EQUIPMENT – PART 1-11: General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment (document IEC/SC62A 62A/624/CD) 2008. |  |  | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [6 empty] | Protect: Depending on the failure, visual and audible alarms appropriate for the condition will alarm. |  | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD257 | [6 empty] | WARNING: To ensure proper servicing and avoid the possibility of physical injury to personnel or damage to the ventilator, only personnel authorized and qualified by Covidien should attempt to service or make authorized modifications to the Puritan Bennett™ 560 Ventilator. Do not attempt to open, repair or otherwise service the ventilator yourself. Doing so might endanger the patient, damage the ventilator, and/or void your warranty. Only personnel authorized and qualified by Covidien should repair, open or service the ventilator. | 2 | 1 | 1 | 2
VH56 | Complete loss of oxygen to the ventilator | Oxygen is not delivered |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD342 (560) |  |  |  |  | The ventilation modes that utilize LOW FIO2 setting shall be CPAP, VOL A/C, PRES A/C, PSV ST, PSIMV and VSIMV Modes. The same adjusted value will be kept in all modes when switching from one to another. |  |  | 2 | 3 | 1 | 1 | P2 scores based on clinical expertise, engineering input, complaint data, and environment of use. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User Error - O2 not connected or obstructed |  | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD343 (560) |  |  |  |  | The LOW FIO2 adjustment range shall be 18 to 90%. |  |  | 2 | 3 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD321 (560) PRD322 (560) | [5 empty] | PB560 CHECK FIO2 SENSOR alarm LOW FIO2 alarm HIGH FIO2 alarm FIO2 SENSOR MISSING alarm CALIBRATE FIO2 alarm | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FIO2 kit) that features a minimum and maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD254 | [6 empty] | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FiO2 kit) that features a minimum and Installation and Assembly maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD421 | [6 empty] | WARNING: The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. Inspect the oxygen coupler (Figure 6-16, item 2) before use to ensure it has its black O-ring attached and in good condition. Do not use an oxygen coupler with a missing, damaged, or worn O-ring. Ensure that the oxygen supply pressure to the machine never exceeds 7 psi (50 kPa) or a flow of 15 lpm. Refer to Table A-8. on page A-4 for volume and sensitivity tolerances. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 1
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH57 | Complete loss of air to the ventilator | Air is not delivered | Blower air inlet blocked. (not sufficient ventilation, ventilation loss) | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD35 | [5 empty] | The air inlet filter is designed such that an unclean state can be easily detected |  | 1 | 3 | 1 | 1 | P2 scores based on clinical expertise including sensitive population of CO2 retainers. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Transport bag incorrectly used. (air entry obstructed, connectors accessible) |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 1 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 1 | 1 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator. | 1 | 1 | 1 | 2
|  |  |  | Dyspnea | Y | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD179 (560) | [5 empty] | LOW VTE Alarm | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 1 | 3 | 1 | 1
[5 empty] | Y | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD 181 | [5 empty] | LOW VTI ALARM |  | 1 | 2 | 1 | 1
[5 empty] | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD442 | [5 empty] | TURB OVERHEAT alarm | WARNING: If the ambient temperature where the device is operated is greater than 35 °C (95 °F), the temperature of the patient circuit or the flow supplied at the device outlet may exceed 41 °C (106 °F), and the patient circuit may reach up to 60°C (140°F). This may lead to undesirable side effects for the patient. To avoid injury to the patient move the patient and the ventilator to a cooler location. For more information, contact Covidien. WARNING: Operating speed of the cooling fan not properly adjusted for the internal ambient temperature of the device. Restart ventilator to see if alarm clears. If not, replace the ventilator and call your customer service representative. | 1 | 1 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD51 |  |  |  |  | Safety Net |  |  | 1 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD53 |  |  |  |  | POST |  |  | 1 | 1 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD200 | [7 empty] | 1 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD442 | [7 empty] | 1 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD229 | [6 empty] | Description of Oxygen function and installation | 1 | 1 | 1 | 2
|  |  |  | Oxygen saturation, High | Y | 2478 | Asymptomatic; no treatment necessary; found on arterial blood gas measurement; may require decreasing supplemental oxygen | 1 | PRD321 (560) PRD322 (560) | [5 empty] | PB560 CHECK FIO2 SENSOR alarm HIGH FIO2 alarm FIO2 SENSOR MISSING alarm CALIBRATE FIO2 alarm | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FIO2 kit) that features a minimum and maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 1 | 2 | 1 | 1
[5 empty] | N | 2478 | Asymptomatic; Oxygen decreased (e.g., for retinal protection-neonate, protocol adherence) | 2 | PRD254 | [6 empty] | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FiO2 kit) that features a minimum and Installation and Assembly maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 1 | 2 | 1 | 1
[6 empty] | 2478 | Symptomatic; Requires elective increase in non-invasive ventilatory support due to decrease of respiratory drive (e.g., COPD) | 3 | [8 empty] | 1 | 2 | 1 | 1
[5 empty] | N | 2478 | Symptomatic; Requires emergent intubation or rescue due to loss of respiratory drive (e.g., COPD, congenital cardiac cyanotic lesions requiring <85% SpO2s) | 4 | PRD357 | [5 empty] | Preventative maintenance schedule - O2 sensor change |  | 1 | 1 | 1 | 2
[5 empty] | N | 2478 | Death | 5 | PRD421 | [6 empty] | WARNING: The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. Inspect the oxygen coupler (Figure 6-16, item 2) before use to ensure it has its black O-ring attached and in good condition. Do not use an oxygen coupler with a missing, damaged, or worn O-ring. Ensure that the oxygen supply pressure to the machine never exceeds 7 psi (50 kPa) or a flow of 15 lpm. Refer to Table A-8. on page A-4 for volume and sensitivity tolerances. | 1 | 1 | 1 | 2
|  |  |  | Atelectasis | N | N/A | Asymptomatic, no intervention | 1 | PRD38 |  |  |  |  | The ventilator provides O2 intake port and connector which allows connecting of low pressure oxygen sources at less than 15 lpm at 50 KPa (e.g. Oxygen concentrator). |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Minor medical intervention indicated | 2 | PRD253 | [6 empty] | WARNING: To avoid injury to the patient and/or possible damage to the ventilator: before using the ventilator, use a flow meter (flow regulator) to regulate the oxygen supply to specifications before connecting the ventilator to the oxygen supply. The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Major medical intervention indicated | 3 | PRD254 | [6 empty] | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FiO2 kit) that features a minimum and Installation and Assembly maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 1 | 1 | 1 | 1
|  |  |  | Tissue Damage/Tissue Trauma (ROP for Neo) | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[5 empty] | N | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Examples: Bleeding vessels requiring suture repair with blood transfusions; additional/unintended tissue resection required to achieve intended patient outcome. Procedure changed to open procedure to manage bleeding. Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 1 | 1 | 1 | 1
[5 empty] | N | 2104 | Sudden significant injury to tissue that requires immediate attention to prevent a catastrophic event; Examples: Profuse bleeding on a large vessel (pulmonary artery/vein, renal artery/vein, SMA/IMA); additional unintended tissue loss/resection needed causing additional medical intervention such as a permanent colostomy. Ventilation: Symptomatic; Signs of ARDS requiring Major change in therapy (e.g., ECMO, HFV). | 4 | PRD431 | [6 empty] | WARNING: Be aware this material describes how to respond to ventilator, but does not tell you how to respond to the patient. | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 3 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH58 | Low inspiratory flow | Flow delivered to the patient is lower than set or expected | User Error - Incorrect settings for prescribed ventilation | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering inputs, and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Inlet filter obstructed |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD35 | [5 empty] | The air inlet filter is designed such that an unclean state can be easily detected |  | 2 | 2 | 1 | 1
|  |  | Transport bag incorrectly used. (air entry obstructed, connectors accessible) |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 2 | 1 | 1
|  |  | DFMEA CPU PCBA - Barometric pressure sensor Inspiratory flow measure Pneumatic fittings - Inspiratory flow sensor fitting |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD231 | [6 empty] | WARNING: The patient circuit is intended for single use by a single patient and should be changed according to the manufacturer’s recommendations and according to the patient circuit lifetime. Refer to the instructions for use supplied by the manufacturer of the patient circuit (included with the ventilator) and Chapter 4, “Installation and Assembly.” | 2 | 1 | 1 | 2
|  |  | SWFMEA Breath Delivery Software target pressures and flows are not aggressive for patient Software target pressures and flows are conservative for patient Software detects an exhalation valve present when there is no physical exhalation valve Exhalation Valve control failure Initiates breaths incorrectly Alarms Incorrect detection of High Peak Pressure No detection of Low Pressure No detection of Occlusion UI - Incorrect Settings change |  | N | 1916 | Death | 5 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
|  |  | Use Error - incorrect set-up of breathing circuit | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator. | 2 | 2 | 1 | 1
|  |  | Use error - failure to change filter |  | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD442 | [5 empty] | TURB OVERHEAT alarm | WARNING: If the ambient temperature where the device is operated is greater than 35 °C (95 °F), the temperature of the patient circuit or the flow supplied at the device outlet may exceed 41 °C (106 °F), and the patient circuit may reach up to 60°C (140°F). This may lead to undesirable side effects for the patient. To avoid injury to the patient move the patient and the ventilator to a cooler location. For more information, contact Covidien. WARNING: Operating speed of the cooling fan not properly adjusted for the internal ambient temperature of the device. Restart ventilator to see if alarm clears. If not, replace the ventilator and call your customer service representative. | 2 | 1 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD246 | [6 empty] | WARNING: Do not start ventilation until you ensure that the device is suitably assembled, that the air inlet filter is properly installed and is not obstructed, and that there is proper clearance all around the unit. Also ensure that the patient circuit is suitably connected to both the ventilator and the patient and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD233 (560) | [6 empty] | WARNING: If exhaled tidal volume measurements are required to ensure correct patient ventilation a double-limb patient circuit configuration must be used in order to detect leaks. In this case, both the minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. Operational checklist - Ensure the patient breathing circuit is correctly attached to the ventilator, with all exhaled volume monitoring is required, use the double-limb circuit for exhaled tidal volume monitoring. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD256 | [6 empty] | WARNING: Avoid using the ventilator, if possible, in dusty environments. Dusty environments may require more vigilant monitoring, cleaning, and/or replacement of air intake and other filters. | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD181 PRD188 | [5 empty] | LOW VTI Alarm |  | 2 | 1 | 1 | 2
[5 empty] | Y | N/A | Death | 5 | PRD179 (560) | [5 empty] | LOW VTE Alarm |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator.
[5 empty] | Y |  |  |  | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms
[5 empty] | Y |  |  |  | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm
[5 empty] | Y |  |  |  | PRD172 | [5 empty] | Apnea alarms
[5 empty] | N |  |  |  | PRD173 | [5 empty] | HIGH RATE alarm
[5 empty] | Y |  |  |  | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms
[5 empty] | N |  |  |  | PRD128 |  |  |  |  | PEEP pressure control accuracy during exhalation steady phase
[5 empty] | N |  |  |  | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold
[5 empty] | Y |  |  |  | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm
[5 empty] | N |  |  |  | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes
[5 empty] | N |  |  |  | PRD113 (560) | [5 empty] | HIGH VTI adjustment range (560)
[5 empty] | N |  |  |  | PRD74 |  |  |  |  | PEEP adjustment range
[5 empty] | N |  |  |  | PRD411 | [6 empty] | WARNING: The level of inspiratory resistance of the circuit and accessories (bacteria filter, humidifier, and so on) must be as low as possible. Settings – particularly the PATIENT DISCONNECTION alarm, High inspired volume (High VTI), and Low inspired volume (Low VTI) settings – must be periodically adjusted according to changes in the patient circuit resistance – especially when filters are replaced.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH59 | Low inspiratory gas pressure | Pressure delivered to the patient is lower than set or expected | User Error - Incorrect settings for prescribed ventilation | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering inputs, and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Inlet filter obstructed |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD35 | [5 empty] | The air inlet filter is designed such that an unclean state can be easily detected |  | 2 | 2 | 1 | 1
|  |  | Transport bag incorrectly used. (air entry obstructed, connectors accessible) |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 2 | 1 | 1
|  |  | DFMEA CPU PCBA - Barometric pressure sensor '+5V Reference Internal pressure sensor measure Pneumatic fittings - Inspiratory flow sensor fitting |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD231 | [6 empty] | WARNING: The patient circuit is intended for single use by a single patient and should be changed according to the manufacturer’s recommendations and according to the patient circuit lifetime. Refer to the instructions for use supplied by the manufacturer of the patient circuit (included with the ventilator) and Chapter 4, “Installation and Assembly.” | 2 | 1 | 1 | 2
|  |  | SWFMEA Breath Delivery Software target pressures and flows are conservative for patient Software detects an exhalation valve present when there is no physical exhalation valve Exhalation Valve control failure Incorrect detection of High Peak Pressure No detection of Low Pressure No detection of Occlusion No detection of "Check Proximal Line 1 " alarm UI - Incorrect Settings change |  | N | 1916 | Death | 5 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
|  |  | Use Error - incorrect set-up of breathing circuit | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator. | 2 | 2 | 1 | 1
|  |  | Use error - failure to change filter |  | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD442 | [5 empty] | TURB OVERHEAT alarm | WARNING: If the ambient temperature where the device is operated is greater than 35 °C (95 °F), the temperature of the patient circuit or the flow supplied at the device outlet may exceed 41 °C (106 °F), and the patient circuit may reach up to 60°C (140°F). This may lead to undesirable side effects for the patient. To avoid injury to the patient move the patient and the ventilator to a cooler location. For more information, contact Covidien. WARNING: Operating speed of the cooling fan not properly adjusted for the internal ambient temperature of the device. Restart ventilator to see if alarm clears. If not, replace the ventilator and call your customer service representative. | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD246 | [6 empty] | WARNING: Do not start ventilation until you ensure that the device is suitably assembled, that the air inlet filter is properly installed and is not obstructed, and that there is proper clearance all around the unit. Also ensure that the patient circuit is suitably connected to both the ventilator and the patient and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD233 (560) | [6 empty] | WARNING: If exhaled tidal volume measurements are required to ensure correct patient ventilation a double-limb patient circuit configuration must be used in order to detect leaks. In this case, both the minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. Operational checklist - Ensure the patient breathing circuit is correctly attached to the ventilator, with all exhaled volume monitoring is required, use the double-limb circuit for exhaled tidal volume monitoring. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD256 | [6 empty] | WARNING: Avoid using the ventilator, if possible, in dusty environments. Dusty environments may require more vigilant monitoring, cleaning, and/or replacement of air intake and other filters. | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD181 PRD188 | [5 empty] | LOW VTI Alarm |  | 2 | 1 | 1 | 2
[5 empty] | Y | N/A | Death | 5 | PRD179 (560) | [5 empty] | LOW VTE Alarm |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator.
[5 empty] | Y |  |  |  | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms
[5 empty] | N |  |  |  | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold
[5 empty] | Y |  |  |  | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm
[5 empty] | N |  |  |  | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes
[5 empty] | N |  |  |  | PRD113 (560) | [5 empty] | HIGH VTI adjustment range (560)
[5 empty] | Y |  |  |  | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms
[5 empty] | N |  |  |  | PRD74 |  |  |  |  | PEEP adjustment range
[5 empty] | Y |  |  |  | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm
[5 empty] | N |  |  |  | PRD173 | [5 empty] | HIGH RATE alarm
[5 empty] | N |  |  |  | PRD411 | [6 empty] | WARNING: The level of inspiratory resistance of the circuit and accessories (bacteria filter, humidifier, and so on) must be as low as possible. Settings – particularly the PATIENT DISCONNECTION alarm, High inspired volume (High VTI), and Low inspired volume (Low VTI) settings – must be periodically adjusted according to changes in the patient circuit resistance – especially when filters are replaced.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH60 | Low inspiratory volume | Volume delivered to the patient is lower than set | User Error - Incorrect settings for prescribed ventilation | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 2 | 2 | 1 | 1 | P2 scores based on clinical expertise, engineering input, complaint data, and environment of use. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Inlet filter obstructed |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD35 | [5 empty] | The air inlet filter is designed such that an unclean state can be easily detected |  | 2 | 2 | 1 | 1
|  |  | Transport bag incorrectly used. (air entry obstructed, connectors accessible) |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 2 | 1 | 1
|  |  | DFMEA CPU PCBA - Barometric pressure sensor Exhalation flow measure Exhalation Valve command Pneumatic fittings - Inspiratory flow sensor fitting |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD231 | [6 empty] | WARNING: The patient circuit is intended for single use by a single patient and should be changed according to the manufacturer’s recommendations and according to the patient circuit lifetime. Refer to the instructions for use supplied by the manufacturer of the patient circuit (included with the ventilator) and Chapter 4, “Installation and Assembly.” | 2 | 1 | 1 | 2
|  |  | SWFMEA Breath Delivery Software target pressures and flows are conservative for patient Software detects an exhalation valve present when there is no physical exhalation valve Exhalation Valve control failure Initiates breaths incorrectly |  | N | 1916 | Death | 5 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
|  |  | ALARMS Incorrect detection of High Peak Pressure No detection of Low Pressure Incorrect monitored exhaled tidal volume (VTE) from flow sensor No detection of Occlusion UI - Incorrect Settings change | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator. | 2 | 2 | 1 | 1
|  |  | Use Error - incorrect set-up of breathing circuit |  | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1
|  |  | Use error - failure to change filter |  | N | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD442 | [5 empty] | TURB OVERHEAT alarm | WARNING: If the ambient temperature where the device is operated is greater than 35 °C (95 °F), the temperature of the patient circuit or the flow supplied at the device outlet may exceed 41 °C (106 °F), and the patient circuit may reach up to 60°C (140°F). This may lead to undesirable side effects for the patient. To avoid injury to the patient move the patient and the ventilator to a cooler location. For more information, contact Covidien. WARNING: Operating speed of the cooling fan not properly adjusted for the internal ambient temperature of the device. Restart ventilator to see if alarm clears. If not, replace the ventilator and call your customer service representative. | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD246 | [6 empty] | WARNING: Do not start ventilation until you ensure that the device is suitably assembled, that the air inlet filter is properly installed and is not obstructed, and that there is proper clearance all around the unit. Also ensure that the patient circuit is suitably connected to both the ventilator and the patient and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD233 (560) | [6 empty] | WARNING: If exhaled tidal volume measurements are required to ensure correct patient ventilation a double-limb patient circuit configuration must be used in order to detect leaks. In this case, both the minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. Operational checklist - Ensure the patient breathing circuit is correctly attached to the ventilator, with all exhaled volume monitoring is required, use the double-limb circuit for exhaled tidal volume monitoring. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD256 | [6 empty] | WARNING: Avoid using the ventilator, if possible, in dusty environments. Dusty environments may require more vigilant monitoring, cleaning, and/or replacement of air intake and other filters. | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD181 PRD188 | [5 empty] | LOW VTI Alarm |  | 2 | 1 | 1 | 2
[5 empty] | Y | N/A | Death | 5 | PRD179 (560) | [5 empty] | LOW VTE Alarm |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator.
[5 empty] | Y |  |  |  | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms
[5 empty] | N |  |  |  | PRD74 |  |  |  |  | PEEP adjustment range
[5 empty] | Y |  |  |  | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms
[5 empty] | N |  |  |  | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold
[5 empty] | Y |  |  |  | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm
[5 empty] | N |  |  |  | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes
[5 empty] | N |  |  |  | PRD128 |  |  |  |  | PEEP pressure control accuracy during exhalation steady phase
[5 empty] | Y |  |  |  | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm
[5 empty] | Y |  |  |  | PRD172 | [5 empty] | Apnea alarms
[5 empty] | N |  |  |  | PRD173 | [5 empty] | HIGH RATE alarm
[5 empty] | N |  |  |  | PRD411 | [6 empty] | WARNING: The level of inspiratory resistance of the circuit and accessories (bacteria filter, humidifier, and so on) must be as low as possible. Settings – particularly the PATIENT DISCONNECTION alarm, High inspired volume (High VTI), and Low inspired volume (Low VTI) settings – must be periodically adjusted according to changes in the patient circuit resistance – especially when filters are replaced.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH61 | Low Peep | PEEP is lower than prescribed setting | User Error - Incorrect settings for prescribed ventilation | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 2 | 2 | 1 | 1 | P2 scores based on clinical expertise including generally lower Oxygen and PEEP requirement for the homecare patient. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA Breath Delivery Software target pressures and flows are conservative for patient Software detects an exhalation valve present when there is no physical exhalation valve Exhalation Valve control failure |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD139 |  |  |  |  | The ventilator provides a numerical display of PEEP for each breath cycle with a display resolution of 1 mbar and within the range 0 to 99 mbar. PEEP value shall be within the current value +/- (2 mbar + 8%). |  |  | 2 | 2 | 1 | 1
|  |  | ALARMS No detection of Occlusion No detection of "Check Proximal Line 1 " alarm User Interface Incorrect Settings change |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD132 |  |  |  |  | The system shall increase the inspiratory pressure of at least 0.1 mbar from the PEEP level within 200 ms after a detectable patient effort. After stabilization of the exhalation bias flow as defined in PRD36, some auto cycling could be tolerated according to the following: OK = Possible to trigger an inspiration with increasing effort from sensitivity setting 0P to 5. AC = Ventilator will have high potential to Auto Cycle C= Lung Compliance |  |  | 2 | 1 | 1 | 1
[5 empty] | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD350 | [5 empty] | Low pressure alarms CHECK EXH VALVE CHECK PROXIMAL LINE1 NO PROXIMAL LINE2 |  | 2 | 1 | 1 | 2
[5 empty] | N | 1916 | Death | 5 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD111 PRD113 (560) | [5 empty] | LOW VTI adjustment range HIGH VTI adjustment range (560) |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD162 | [5 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
|  |  |  | Hyperventilation | Y | 1910 | Asymptomatic; PaCO2 or ETCO2 low. Mild tachypnea. | 1 | PRD103 PRD104 PRD105 PRD106 | [5 empty] | LOW PRESSURE alarms |  | 2 | 2 | 1 | 1
[5 empty] | N | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD128 |  |  |  |  | PEEP pressure control accuracy during exhalation steady phase |  |  | 2 | 2 | 1 | 1
[5 empty] | Y | 1910 | Symptomatic; PaCO2 or ETCO2 low. Tachypnea/↑ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 1910 | Symptomatic; PaCO2 or ETCO2 unacceptable range, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms |  | 2 | 1 | 1 | 2
[5 empty] | N | 1910 | Death | 5 | PRD74 |  |  |  |  | PEEP adjustment range |  |  | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH62 | Low O2 delivery | Device delivers O2 below the prescribed oxygen setting | User Error - Incorrect settings for prescribed ventilation | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD254 | [6 empty] | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FiO2 kit) that features a minimum and Installation and Assembly maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 1 | 2 | 1 | 1 | P2 scores based on clinical expertise including generally lower Oxygen requirement for the homecare patient. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA: FIO2 monitoring Valve O2 connector O2 valve command |  | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD229 | [6 empty] | Description of Oxygen function and installation | 1 | 2 | 1 | 1
|  |  | DFMEA O2 SUPPLY: O2 Solenoid Valve tubings and connectors oxygen connector - Connecting to the ventilator |  | Y | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD321 (560) PRD322 (560) | [5 empty] | PB560 CHECK FIO2 SENSOR alarm LOW FIO2 alarm HIGH FIO2 alarm FIO2 SENSOR MISSING alarm CALIBRATE FIO2 alarm | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FIO2 kit) that features a minimum and maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 1 | 2 | 1 | 1
|  |  | SWFMEA Monitors - No detection of FIO2 alarm condition |  | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD253 | [6 empty] | WARNING: To avoid injury to the patient and/or possible damage to the ventilator: before using the ventilator, use a flow meter (flow regulator) to regulate the oxygen supply to specifications before connecting the ventilator to the oxygen supply. The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. | 1 | 1 | 1 | 2
|  |  | Misidentififcation of oxygen inlet connector |  |  | 2477 | Death | 5 | PRD38 | [7 empty] | 1 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD421 | [6 empty] | WARNING: The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. Inspect the oxygen coupler (Figure 6-16, item 2) before use to ensure it has its black O-ring attached and in good condition. Do not use an oxygen coupler with a missing, damaged, or worn O-ring. Ensure that the oxygen supply pressure to the machine never exceeds 7 psi (50 kPa) or a flow of 15 lpm. Refer to Table A-8. on page A-4 for volume and sensitivity tolerances. | 1 | 2 | 1 | 1
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 1 | 2 | 1 | 1
[5 empty] | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD431 | [6 empty] | WARNING: Be aware this material describes how to respond to ventilator, but does not tell you how to respond to the patient. | 1 | 2 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 1 | 1 | 2
[5 empty] | N | 1816 | Death | 5 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD162 | [5 empty] | graphics to display patient response |  | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD253 | [6 empty] | PB560 WARNING: Strictly follow the instructions provided in section 4.8.2, “Connecting the Oxygen Supply,” which include the use of a flow regulator and special oxygen connector. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD420 | [6 empty] | PB560 WARNING: The Puritan Bennett™ 560 Ventilator is designed to deliver a percentage of oxygen equal or lower than 50%. Do not exceed this value as this may cause the ventilator to malfunction and put the patient at risk. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD421 | [6 empty] | PB560 WARNING: Ensure that the only gas supplied to the ventilator through the dedicated oxygen supply connector is medical-grade oxygen. | 1 | 2 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD421 | [6 empty] | PB560 WARNING: The oxygen supply hose ages even when it is not in use and should be replaced periodically. Follow the expiration date, if any. | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD421 | [6 empty] | PB560 WARNING:Before connecting the oxygen supply, ensure that the stud on the oxygen connector is protruding outwards.
VH63 | O2 sensor failure | Failure impacts calibrations and/or monitored oxygen values(fi02) | incorrect sensor fitted to the ventilator - range of the sensor is not calibrated to the ventilator | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD339 |  |  |  |  | Patient circuit configuration may be selected prior to calibration |  | Note: Calibration may be done with either an adult or pediatric circuit; however, Pediatric YES/NO must be appropriately selected in the ventilator Preferences Menu. | 2 | 2 | 1 | 1 | P2 scores based on clinical expertise including generally lower Oxygen requirement for the homecare patient. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA FIO2 monitoring FIO2 connector |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD357 | [5 empty] | Preventative maintenance schedule - O2 sensor change |  | 2 | 2 | 1 | 1
|  |  | PM schedule not adhered to/ Caliration not performed |  | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD321 (560) PRD322 (560) | [5 empty] | PB560 CHECK FIO2 SENSOR alarm LOW FIO2 alarm HIGH FIO2 alarm FIO2 SENSOR MISSING alarm CALIBRATE FIO2 alarm | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FIO2 kit) that features a minimum and maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 2 | 2 | 1 | 1
|  |  |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD253 | [6 empty] | WARNING: To avoid injury to the patient and/or possible damage to the ventilator: before using the ventilator, use a flow meter (flow regulator) to regulate the oxygen supply to specifications before connecting the ventilator to the oxygen supply. The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. | 2 | 1 | 1 | 1
[5 empty] | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [7 empty] | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FIO2 kit) that features a minimum and maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 2 | 1 | 1 | 1
|  |  |  | Hypoventilation (COPD) | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 2 | 1 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD431 |  |  |  |  | The ventilator user and clinician manuals shall include a warning informing the user that the manual describes how to respond to ventilator alarms, but it does NOT tell you how to respond to the patient. |  | WARNING: Be aware this material describes how to respond to ventilator, but does not tell you how to respond to the patient. | 2 | 1 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
VH64 | Incorrect filter used | External filter (Insp or Exp) filter media does not meet specification. | User Error - incorrect filter fitted to the ventilator filter media allows greater particulate matter into the ventilator than the approved filter | Dyspnea | Y | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical experience, complaint data and engineering expertise. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User Error - unapproved filter fitted to the ventilator filter media is closer woven leading to high pressure drop across the filter |  | Y | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 2 | 2 | 1 | 1
|  |  | Filter applied in reverse orientation than intended (inidicated on filter)/in the incorrect location |  | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD180 | [5 empty] | The ventilator shall initiate a high priority HIGH VTI alarm when the VTI measured value is above the HIGH VTI setting for the following ventilation modes CPAP, PRES A/C, PSV ST, PSIMV. HIGH VTI Alarm - high priority alarm - Inspired tidal volume greater than Max VTI set during three consecutive breaths in PSV, CPAP, P A/C, P SIMV, and V SIMV modes. Alarm activation occurs: After three consecutive breaths. |  | 2 | 1 | 1 | 2
[5 empty] | Y | 1816 | Death | 5 | PRD 180 PRD181 PRD188 | [5 empty] | HIGH VTI ALARM LOW VTI Alarm |  | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD230 | [6 empty] | Description of patient circuit | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD273 | [5 empty] | All device operational connectors and ports are labeled |  | 2 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 2 | 1 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD234 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD256 | [6 empty] | WARNING: Avoid using the ventilator, if possible, in dusty environments. Dusty environments may require more vigilant monitoring, cleaning, and/or replacement of air intake and other filters. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD17 |  |  |  |  | The ventilator testing demonstrates the gas pathway materials biocompatibility applying testing requirements from per Environmental Protection Agency (EPA) method TO 15 Determination Of Volatile Organic Compounds (VOCs) In Air Collected In Specially-Prepared Canisters And Analyzed By Gas Chromatography/Mass spectrometry (GC/MS). Any volatile organic compound levels detected shall be less than 10% of the 8 hour US OSHA Permissible Exposure Limit (PEL) for General Industry as indicated in 29 CFR 1910 Subpart Z. The ventilator testing shall also demonstrate that the weights of particulate which size are 2.5 micron and below, is less that 15 micrograms/cubic meter. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD35 | [5 empty] | The air inlet filter is designed such that an unclean state can be easily detected |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD230 | [6 empty] | Description of patient circuit | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 1 | 1 | 2
|  |  |  | Infection | N | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth or inhalation, Incision/drainage see abscess | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | 1930 | Symptomatic; IV antibiotics required with increased ventilatory support. An infection that requires more advanced medical intervention | 3 | [5 empty] | PB560 Ventilator is designed to meet the 5356-1 (Anesthetic and respiratory equipment - Conical connectors). |  |  | 2 | 2 | 1 | 1
[6 empty] | 1930 | Symptomatic; IV antibiotics required, increased ventilatory support, isolation, refractory to treatment. | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | 1930 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 2 | 1 | 1
|  |  |  | Foreign body in patient |  | 2687 | Minor Inflammation, no treatment needed. Asymptomatic Ventilation: and/or patient clears on their own. | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2687 | Symptomatic; requires minor treatment such as oral medication for symptoms. | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2687 | Altered function present (i.e. pain/infection); requiring hospitalization and potential for elective non-urgent operative intervention for surgical repair. Ventilation: Symptomatic; requires suction, Chest PT. | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 2687 | Life threatening consequences; resulting in signs symptoms of shock or obstruction; removal of foreign body needed immediately to repair damage. (i.e. Fresh bullet for gun shot near critical structure, obstruction of airway resulting in difficult breathing requiring intubation, etc.). Emergent invasive intervention required, e.g. rigid or flexible bronchoscopy. | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | 2687 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH65 | Excessive Light | Bright light from display/LEDs. Inability to dim the screen (e.g. during sleeping hours) | User Error - preferences not set | anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD28 |  |  |  |  | The ventilator provides a means to manually adjust the contrast level of the LCD. | - |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD29 |  |  |  |  | The ventilator display system includes a screen saver function with the possibility to be disabled by the operator. |  |  | 1 | 2 | 1 | 1
|  |  |  | Fatigue | N | 1849 | Minor fatigue that comes and goes | 1 | PRD369 | [5 empty] | Two consecutive presses on the AUDIO PAUSE / ALARM PAUSE key cancel all cancelable active alarms. This includes auditory signal, message display and Led signals. |  | 1 | 2 | 1 | 1
[5 empty] | N | 1849 | Symptomatic; fatigue lasting longer than a few hours or days unrelieved by rest. | 2 | - | [6 empty] | The screen backlight never switches off during ventilation Backlight set to YES in Preferences menu. Set Backlight to OFF in Preferences menu. | 1 | 2 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH66 | Missing filter | External filter (Insp or Exp) is missing from ventilator breathing circuit | User Error - filter not fitted/ incorrectly fitted | Infection | N | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | PRD230 | [6 empty] | Description of patient circuit | 2 | 2 | 1 | 1 | P2 scores are based on clinical experience, literature. and complaint data information. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA FILTRATION SYSTEM - Air inlet filter |  | N | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth or inhalation, Incision/drainage see abscess | 2 | PRD240 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. | 2 | 2 | 1 | 1
[5 empty] | N | 1930 | Symptomatic; IV antibiotics required with increased ventilatory support. An infection that requires more advanced medical intervention | 3 | PRD273 | [5 empty] | All device operational connectors and ports are labeled |  | 2 | 2 | 1 | 1
[5 empty] | N | 1930 | Symptomatic; IV antibiotics required, increased ventilatory support, isolation, refractory to treatment. | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 2
[5 empty] | N | 1930 | Death | 5 | PRD262 | [6 empty] | Manual Instruction-It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. In a double-limb configuration, two bacteria filters are used: one at | 2 | 2 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 2 | 1 | 1
|  |  |  | Cross Patient Exposure To Body Fluids |  | 1745 | No medical treatment required if in contact with in-tack skin | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1745 | prophalacric treatment required for exposure to blood or body fluid | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1745 | prophalacric treatment required for exposure to blood or body fluid | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1745 | life threatening infection, hospitalization, or prolonged hospitalization, or greater than 3 drugs | 4 | [8 empty] | 2 | 2 | 1 | 2
[6 empty] | 1745 | Death | 5 | [8 empty] | 2 | 2 | 1 | 2
VH67 | Use error of Respiratory Insight software | User has an incorrect knowledge about the software abilities and intended use. | Respiratory Insight software Patient Data not protected | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD391 |  |  |  |  | Design Software Software database reports issued do not incorporate the patient name but only a patient ID |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical and engineering expertise. Respiratory insight software is a service tool. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Respiratory Insight software Inadvertence data delete | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD402 |  |  |  |  | Design Software A warning message in Respiratory Insight software is displayed when the user tries to delete data |  |  | 1 | 2 | 1 | 1
|  |  | Respiratory Insight software Recording data error |  | N |  |  |  | PRD401 |  |  |  |  | Design Software A warning message in the Respiratory Insight software is displayed when the user tries recording the same data file on two different patient profile and form associating one device to several patients
|  |  | Misindetification of USB Connection Cable |  | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD370 | [6 empty] | Respiratory Insight Software manual WARNING: Before using the software and associated equipment, read, understand, and strictly follow the information in Chapter 1, “Introduction”. WARNING: The Puritan Bennett™ Respiratory Insight Software is intended to support the display and analysis of ventilator data only and is not a diagnostic tool. WARNING: Deletion erases ALL files present on the USB memory device.
VH68 | Blockage of external Filter(s) | Obstructed (Insp or Exp) filter leading to increase or decrease in pressure, flow, or volume | Dirty filter: inspiration filter, exhalation filter, inlet filter. (insufficient ventilation) | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD35 | [5 empty] | The air inlet filter is designed such that an unclean state can be easily detected |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, literature, and engineering inputs. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA FILTRATION SYSTEM - Air inlet filter |  | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 2 | 1 | 1
|  |  | Use Error - addition of multiple filters in place of HME in a single limb circuit |  | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD264 | [5 empty] | PM Schedule | Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 2 | 1 | 1
|  |  | Use error - failure to change filter |  | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD251 | [6 empty] | WARNING: Regularly check the cleanliness of the air inlet filter located on the rear of the ventilator. If necessary, replace the filter before the recommended replacement period is over (see Chapter 8, “Routine Maintenance”). This is particularly important when the ventilator is installed on a wheelchair, because environmental conditions may cause the filter to become dirty more rapidly. WARNING: The air inlet filter is not reusable; do not attempt to wash, clean, or reuse it. Failing to replace a dirty air inlet filter, or operating the ventilator without a filter, may cause serious damage to the ventilator. | 2 | 1 | 1 | 2
[5 empty] | Y | 1816 | Death | 5 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD442 | [5 empty] | TURB OVERHEAT alarm | WARNING: If the ambient temperature where the device is operated is greater than 35 °C (95 °F), the temperature of the patient circuit or the flow supplied at the device outlet may exceed 41 °C (106 °F), and the patient circuit may reach up to 60°C (140°F). This may lead to undesirable side effects for the patient. To avoid injury to the patient move the patient and the ventilator to a cooler location. For more information, contact Covidien. WARNING: Operating speed of the cooling fan not properly adjusted for the internal ambient temperature of the device. Restart ventilator to see if alarm clears. If not, replace the ventilator and call your customer service representative. | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD246 | [6 empty] | WARNING: Do not start ventilation until you ensure that the device is suitably assembled, that the air inlet filter is properly installed and is not obstructed, and that there is proper clearance all around the unit. Also ensure that the patient circuit is suitably connected to both the ventilator and the patient and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD233 (560) | [6 empty] | WARNING: If exhaled tidal volume measurements are required to ensure correct patient ventilation a double-limb patient circuit configuration must be used in order to detect leaks. In this case, both the minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. Operational checklist - Ensure the patient breathing circuit is correctly attached to the ventilator, with all exhaled volume monitoring is required, use the double-limb circuit for exhaled tidal volume monitoring. | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD256 | [6 empty] | WARNING: Avoid using the ventilator, if possible, in dusty environments. Dusty environments may require more vigilant monitoring, cleaning, and/or replacement of air intake and other filters. | 2 | 1 | 1 | 2
|  |  |  | Pneumothorax | N | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 2012 | Symptomatic intervention indicated (e.g. tube placement without sclerosis (no pleurodesis), hospitalization observation, electively place chest tube or | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[6 empty] | 2012 | Sclerosis and/or operative intervention indicated; hospitalization indicated (Pleurodesis) Ventilation: Pneumothorax requiring placement chest tube with hospitalization or prolonged hospitalization | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 2012 | Life-threatening consequences; Tension pneumothorax; airway compromise requiring intubation; urgent surgical intervention or emergency Chest tube placement indicated | 4 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH69 | Inadequate ventilator maintenance | Machine not properly serviced. | User Error - schedule not adhered to | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD356 | [6 empty] | WARNING: To ensure proper performance of the ventilator, the preventative maintenance schedule should be followed. For further information contact Covidien. | 2 | 2 | 1 | 1 | P2 scores are based on Clinical expertise, engineering inputs and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | [6 empty] | PM Schedule |  | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD264 | [6 empty] | Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD296 | [6 empty] | Manual - Operational Verification Checklist | 2 | 1 | 1 | 2
[5 empty] | N | 1816 | Death | 5 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH70 | Use error of proximal flow sensor | Inaccurate display or no display of proximal flow sensor values. | N/A - The PB500 series ventilator proximal flow sensor is changed by service personnel
VH71 | Display field(s) are not visible and/or erratic | Patient or device (monitored) displayed data is not visible or readable to user/clinician | USER INTERFACE LCD does not update or incorrect data is displayed | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | [6 empty] | Protective - Audio alarms will continue to annunciate |  | 1 | 2 | 1 | 1 | P2 scores are based on Clinical expertise, engineering inputs and complaint data. Ventilator transfer is likely result as ventilation continues. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | difficulty reading display while ventilator is within travel bag |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD50 | [5 empty] | POST |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD32 | [5 empty] | Protective - alarm volume is adjustable 65 to 80 dBA through the preference menu |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD296 | - | - | - | - |  |  | Manual - Operational Verification Checklist - Verify the alarm volume is adapted to the patient environment. | 1 | 3 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 1 | 3 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 1 | 3 | 1 | 2
VH72 | No Hazard | No hazard occurred during reported event or identified after event analysis | DFMEA - CPU PCBA Microcontroller +5Vref Remote alarm or nurse call Monitoring memory uart extension Supply +5V '+10V Reference Power Supply Board connector keyboard connector O2 flow sensor connector (not installed on PCBA) SPO2 connector | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1 | P2 scores are based on Clinical expertise, engineering inputs and complaint data. Ventilator transfer is likely result as ventilation continues. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA - Power Supply PCBA Supply switch 24 V Management (Step Up/Down Voltage transformer) Current control for battery charge (Constant current Generator) reference 3 V supply Micro-controller Micro-controller connector USB B (Convertor USB /2UART) USB B connector CPU connector DC external supply (Alimentation DC extern) BATTERY CONNECTOR CONTACTS Turbine supply (turbine board supply) Cooling fan connector cooling fan control (ventilator) microcontroller connector | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
|  |  | DFMEA Buzzer PCBA DC/DC (5V to 12V) Convertor Supply Switch (5V; battery) Power supply loss management Buzzer PCBA back up coin battery Buzzer Coin Battery charger Oscillator Buzzer amplifier Buzzer watch function CPU/ Buzzer Board connector Shutdown DC/DC management |  |  | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
|  |  | DFMEA KEYBOARD MEMBRANE - user interface POWER SUPPLY PCBA blower control cooling fan control (ventilator) |  |  | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  | DFMEA BATTERY CONNECTION PCBA - filtering ENCLOSURES lower enclosure upper enclosure LABELS - cannot read information TURBINE CONTROL PCBA blower control POWER SUPPLY - internal battery |  |  | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
|  |  | User unable to switch vent mode into standby |  |  | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH73 | Noise | Device or device accessory emits loud audible noise | Ventilator's mechanical noise | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD25 |  |  |  |  | During normal usage, the noise level shall not exceed 30 dBA + 10% sound according to ISO 17510-1 2007 standard conditions. |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical experience, engineering expertise and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Alarm volume |  | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD6 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment - Part 1: General Requirements for Safety -2- Collateral standard: Usability IEC 60601-1-6, EN 60601-1-6 |  |  | 1 | 2 | 1 | 1
|  |  | DFMEA CPU PCBA - Turbine connector Blower - Turbine - Soundproofing Enclosures BLOWER enclosure silicone feet |  | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD7 |  |  |  |  | PRD7 The ventilator shall comply with General Requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems IEC 60601-1-8, EN 60601-1-8. |  |  | 1 | 2 | 1 | 1
|  |  | DFMEA POWER SUPPLY PCBA cooling fan control (ventilator) |  | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | PRD417 |  |  |  |  | The ventilator user and clinician manuals shall include a warning instructing the user to adjust the alarm volume to an appropriate level and to be careful not to obstruct the buzzers outlet. |  | Manual - WARNING: Alarm volume should be adjusted with respect to the ventilator’s operating environment and so that the patient's caretakers can hear the alarms. The audible alarm vents located at the front of the device should never be obstructed. The alarm can be paused with the Alarm Pause function by pressing the ALARM CONTROL key twice once the alarm has been declared. | 1 | 2 | 1 | 2
[5 empty] | N | 2328 | See cause of anxiety | 5 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 2 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 2 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 2 | 1 | 1
VH 74 | Deleted
VH75 | Blockage in ventilator system and/or breathing system | Obstruction leading to changes in ventilatory delivery. | Foreign material in patient circuit airway, not sufficient ventilation. | Oxygen saturation, Low (Hypoxemia) | Y | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms |  | 2 | 2 | 1 | 1 | P2 scores based on clinical expertise, literature, engineering inputs, and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA O2 SUPPLY - tubings and connectors |  | Y | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD179 (560) | [5 empty] | LOW VTE Alarm |  | 2 | 2 | 1 | 1
|  |  | DFMEA INSPIRATORY BLOCK - Screen EXHALTION FLOW SYSTEM - Exhalation flow tube |  | Y | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1
|  |  | Misidentififcation of oxygen inlet connector |  | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD233 (560) | [6 empty] | WARNING: If exhaled tidal volume measurements are required to ensure correct patient ventilation a double-limb patient circuit configuration must be used in order to detect leaks. In this case, both the minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. Operational checklist - Ensure the patient breathing circuit is correctly attached to the ventilator, with all exhaled volume monitoring is required, use the double-limb circuit for exhaled tidal volume monitoring. | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD234 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. | 2 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD230 | [6 empty] | Description of patient circuit | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
[5 empty] | N | 1816 | Death | 5 | PRD233 (560) | [6 empty] | WARNING: If exhaled tidal volume measurements are required to ensure correct patient ventilation a double-limb patient circuit configuration must be used in order to detect leaks. In this case, both the minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. Operational checklist - Ensure the patient breathing circuit is correctly attached to the ventilator, with all exhaled volume monitoring is required, use the double-limb circuit for exhaled tidal volume monitoring. | 2 | 1 | 1 | 2
|  |  |  | Pneumothorax | N | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 2012 | Symptomatic intervention indicated (e.g. tube placement without sclerosis (no pleurodesis), hospitalization observation, electively place chest tube or | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | 2012 | Sclerosis and/or operative intervention indicated; hospitalization indicated (Pleurodesis) Ventilation: Pneumothorax requiring placement chest tube with hospitalization or prolonged hospitalization | 3 | PRD262 | [6 empty] | Manual - It is highly recommended that you install a bacteria filter (see Figure 4-12. ) on both single and double-limb circuits. | 2 | 2 | 1 | 1
[5 empty] | Y | 2012 | Life-threatening consequences; Tension pneumothorax; airway compromise requiring intubation; urgent surgical intervention or emergency Chest tube placement indicated | 4 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH76 | Loose External mechanical interface or connection | Any external parts, connectors, or components that become that become loose. | User Error - Physical damage from impact or vibration, or transport bag fault | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the relevant clauses of IEC 60601-1: • Class II Equipment • Internally Powered Equipment • Detachable power supply cord |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Service Error |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD8 |  |  |  |  | The PB560 shall comply with Lung Ventilators for Medical Use- Particular Requirements for Basic Safety and Essential Performance Part 2: Home Care Ventilators for Ventilator-Dependent Patient EN ISO 10651-2 |  |  | 2 | 2 | 1 | 1
|  |  | DFMEA FILTRATION SYSTEM Low inspiratory flow Inspiratory bacteria filter Exhalation Bacteria filter |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD255 | [6 empty] | WARNING: To minimize the risk of damage, you must use the Dual Bag to transport the Puritan Bennett™ 560 Ventilator. See Figure D-2. Dual Bag. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD314 |  |  |  |  | The ventilator shall comply with sections 21.6a (vibration sinusoidal), 21.6b (random vibration) and 21.6c (bump test) of Lung ventilator for medical use - Part 3: Particular requirements for emergency transport ventilators ISO 10651-3. |  |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD364 |  |  |  |  | The ventilator shall be compatible with drop test according to 10.1.3 d article of IEC 60601-1-11 MEDICAL ELECTRICAL EQUIPMENT – PART 1-11: General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment (document IEC/SC62A 62A/624/CD) 2008. |  |  | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [6 empty] | Protect: Depending on the failure, visual and audible alarms appropriate for the condition will alarm. |  | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | [5 empty] | PB560 Ventilator is designed to meet the 5356-1 (Anesthetic and respiratory equipment - Conical connectors). |  |  | 2 | 1 | 1 | 2
VH77 | Patient data/logs not recorded | Ventilator does not record patient information /date time of incidents /inability to download information | Inadvertent disconnection of USB stick | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD264 | [5 empty] | PM Schedule | Service Manual - WARNING - Do not start ventilation until you ensure the device is correctly assembled, that the air inlet filter is properly installed and unobstructed, and that there is proper clearance all around the unit. Also ensure the patient circuit is correctly connected to both the ventilator and the patient, and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use of a USB stick with inappropriate format (not FAT 32) |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD257 | [6 empty] | WARNING: To ensure proper servicing and avoid the possibility of physical injury to personnel or damage to the ventilator, only personnel authorized and qualified by Covidien should attempt to service or make authorized modifications to the Puritan Bennett™ 560 Ventilator. Do not attempt to open, repair or otherwise service the ventilator yourself. Doing so might endanger the patient, damage the ventilator, and/or void your warranty. Only personnel authorized and qualified by Covidien should repair, open or service the ventilator. | 2 | 2 | 1 | 1
|  |  | DFMEA CPU PCBA: Serial Bus Clock events recording Power Supply Board connector |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD401 |  |  |  |  | The clinical software includes a warning popup message to prevent the user recording the same data file on two different patient |  |  | 2 | 2 | 1 | 1
|  |  | DFMEA Power Supply PCBA: 3.3 V supply to USB USB B (Convertor USB /2UART) CPU connector USB HOST USB HOST Connector | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD402 |  |  |  |  | The clinical software includes a warning popup message to prevent the user from deleting data. |  |  | 2 | 3 | 1 | 1
|  |  | DFMEA USB/PC ports and cable - Connection to ventilator |  | N |  |  |  | PRD358 | [6 empty] | WARNING: Always verify the file ID before using a USB memory device to transfer data between the ventilator and a PC.
|  |  | SWFMEA Respiratory Insight Software Failed to record data correctly Inadvertence data loss |  | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N | [10 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH78 | Ventilator and/or accessories tips, moves or falls | The ventilator or ventilator assemblies moves, falls or tips on Patient, user, bystander. Impact on environment also possible. | Use Error - ventilator placed in an inappropriate position | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PRD314 |  |  |  |  | The ventilator shall comply with sections 21.6a (vibration sinusoidal), 21.6b (random vibration) and 21.6c (bump test) of Lung ventilator for medical use - Part 3: Particular requirements for emergency transport ventilators ISO 10651-3. | - |  | 3 | 3 | 2 | 1 | P2 scores are based on clinical experience and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Cart wheel falls off, uneven floor or roll over something that causes imbalance |  | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | PRD364 |  |  |  |  | The ventilator shall be compatible with drop test according to 10.1.3 d article of IEC 60601-1-11 MEDICAL ELECTRICAL EQUIPMENT – PART 1-11: General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment (document IEC/SC62A 62A/624/CD) 2008. |  |  | 3 | 2 | 1 | 1
|  |  | Dual or carrying bag strap fails |  | N | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Examples: Bleeding vessels requiring suture repair with blood transfusions; additional/unintended tissue resection required to achieve intended patient outcome. Procedure changed to open procedure to manage bleeding. Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | PRD18 |  |  |  |  | The base model ventilator weight shall be less than 5 kg. |  |  | 3 | 2 | 1 | 1
|  |  | Failure to use travel bag and insecure placement of ventilator | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD255 | [6 empty] | WARNING: To minimize the risk of damage, you must use the Dual Bag to transport the Puritan Bennett™ 560 Ventilator. See Figure D-2. Dual Bag. | 3 | 2 | 1 | 1
|  |  | Misidentification of the ventilator cart |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD20 |  |  |  |  | The ventilator is designed such that it can be placed at the bedside (freestanding on a nightstand). To do so, the dimensions of the ventilator shall be no more than 250 x 350 mm on the horizontal plane. |  |  | 3 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD21 |  |  |  |  | The ventilator is designed such that it can be securely mounted on a wheelchair or walker. |  |  | 3 | 2 | 1 | 1
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD446 | [6 empty] | WARNING: To ensure stability, when the Puritan Bennett™ 560 Ventilator is mounted on a cart, the weight of the oxygen bottle should not exceed 14 kg (30 lbs.). | 3 | 2 | 1 | 1
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | [7 empty] | WARNING: Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. | 3 | 2 | 1 | 1
[5 empty] | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 3 | 2 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD451 | - | - | - | - | A cart with the following characteristics is provided as an accessory: • The cart shall meet the rough handling, stability, and suspension systems tests per UL 60601-1 with the ventilator attached. • The cart shall have a pole mount able to accommodate a basket, humidifier bracket and oxygen cylinder. • The cart shall be able to accommodate a circuit arm. • Cart shall be designed to accommodate a cylinder in the range of the following configurations: B, C, D, E and G-cylinders (O2 and air), with 6.89 inches (17.5 cm) maximum diameter, maximum 30.31 inches (77 cm) height, and a minimum 10 inches (25.5 cm) height |  |  | 3 | 1 | 1 | 2
[5 empty] | N | 1816 | Death | 5 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 3 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 3 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 3 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD247 | [6 empty] | PB560 - WARNING: Place the ventilator in a safe place when ventilating and according to the recommendations in the manual. | 3 | 1 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 3 | 1 | 1 | 2
|  |  |  | Anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [8 empty] | 3 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 3 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 3 | 1 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 3 | 1 | 1 | 2
|  |  |  | Delay of Treatment |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 3 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 3 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 3 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 3 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 3 | 1 | 1 | 2
VH79 | Pressure line (including prox flow tubing) disconnect | Reduced ventilator support | User Error - Prox line is not connected or is pulled loose User Error - incorrect patient circuit used | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD350 | [6 empty] | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 2 | 3 | 1 | 1 | P2 scores are based on clinical experience and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA Alarms No detection of "Check Proximal Line 1 " alarm |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD273 | [5 empty] | All device operational connectors and ports are labeled |  | 2 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD248 | [6 empty] | WARNING: The patient circuit should always be positioned to avoid hindering the patient's movements, to prevent accidental disconnection or leakage, and to minimize the risk of patient strangulation. | 2 | 1 | 1 | 2
[5 empty] | N | 1916 | Death | 5 | PRD189 (560) | [5 empty] | HIGH RATE alarm (due to autotriggering) |  | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | Y | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 2 | 2 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD 180 PRD181 PRD188 | [5 empty] | HIGH VTI ALARM LOW VTI Alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD103 PRD104 PRD105 PRD106 | - | - | - | - |  | LOW PRESSURE alarms |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 1 | 1 | 2
[5 empty] | Y | 2477 | Death | 5 | PRD110 PRD111 | - | - | - | - |  | low inspiratory tidal volume alarm LOW VTI alarm |  | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD320 | [5 empty] | Leakage alarms HIGH LEAKAGE Alarm EXH VALVE LEAKAGE ALARM E SENS FAULT OR CIRC LEAK |  | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD162 | - | - | - | - |  | Graphics to display patient response - |  | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD240 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD306 | [6 empty] | Manual - WARNING: The patient circuit should not be changed during ventilation. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD415 | [6 empty] | Manual - WARNING: When using non-invasive ventilation (NIV), without an exhalation valve, use a vented nose or face mask or a non vented combined with a leak accessory. When using noninvasive ventilation (NIV), with an exhalation valve, use a non vented mask. | 2 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
VH80 | Pressure line (including prox flow tubing) is kinked or obstructed | Impact on ventilatory support due to changes in pressure, flow and/or volume | User Error - Exhalation valve, proximal pressure line reversed. (not sufficient ventilation or ventilation loss) | Hypoventilation | Y | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering expertise and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use Error- pressure line kinked during set up or during patient movement |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 2 | 2 | 1 | 1
|  |  | SWFMEA Alarms No detection of "Check Proximal Line 1 " alarm |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD 180 PRD181 PRD188 | [5 empty] | HIGH VTI ALARM LOW VTI Alarm |  | 2 | 1 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD234 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. | 2 | 1 | 1 | 2
[5 empty] | Y | 1916 | Death | 5 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 2 | 1 | 1 | 2
|  |  |  | Pneumothorax | Y | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated Asymptomatic-clinical or diagnostic finding. | 1 | PRD103 PRD104 PRD105 PRD106 | - | - | - | - |  | LOW PRESSURE alarms |  | 2 | 2 | 1 | 1
[5 empty] | N | 2012 | Small to Moderate Pneumothorax, electively place chest tube or observation | 2 | PRD107 PRD108 PRD109 | [5 empty] | High pressure threshold |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2012 | Large Pneumothorax, electively place chest tube | 3 | PRD110 PRD111 | - | - | - | - |  | low inspiratory tidal volume alarm LOW VTI alarm |  | 2 | 1 | 1 | 1
[5 empty] | N | 2012 | Tension Pneumothorax; Emergent Chest tube placement, intubation. | 4 | PRD112 | [5 empty] | High VTI settings in CPAP, PRES A/C, PSV ST, and PSIMV Modes |  | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD162 | - | - | - | - |  | Graphics to display patient response - |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD230 | [6 empty] | Description of patient circuit | 2 | 2 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD190 | [5 empty] | OCCLUSION CHECK CIRCUIT alarms |  | 2 | 1 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD260 | [6 empty] | WARNING: Before opening the packaging for the Patient Circuit, ensure that no damage is evident to the packaging or its contents. Do not use if evidence of damage exists. | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD248 | [6 empty] | Manual WARNING: The patient circuit should always be positioned to avoid hindering the patient's movements, to prevent accidental disconnection or leakage, and to minimize the risk of patient strangulation. | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH81 | EMC Susceptibility | Interference of the ventilators ability to provide ventilation due to electro magnetic energy. | Electromagnetic interference | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD4 |  |  |  |  | The ventilator complies with Medical Electrical Equipment - Part 1: General Requirements for Safety -2- Collateral standard Electro-Magnetic Compatibility requirements and tests IEC 60601-1-2, EN 60601-1-2, excluding the marking requirements for Essential Performance. When used with USB cable (mini connector), the ventilator shall comply with Class A radiated emissions |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering expertise and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Electromagnetic interference, failure in detecting bad ventilation |  | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD8 |  |  |  |  | The ventilator complies with Lung Ventilators for Medical Use- Particular Requirements for Basic Safety and Essential Performance Part 2: Home Care Ventilators for Ventilator-Dependent Patient EN ISO 10651-2. |  |  | 1 | 2 | 1 | 1
|  |  | Electromagnetic interference, Ventilator shut down |  | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD335 |  |  |  |  | The ventilator complies with section 21 (emission of Radio Frequency Energy) of Environmental Conditions and Test Procedures for Airborne Equipment, RTCA DO-160E 2007. |  |  | 1 | 1 | 1 | 2
|  |  | Excessive radio frequency radiation |  | N | 2484 | Death | 5 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 1 | 1 | 1 | 2
|  |  | DFMEA CPU PCBA - Microcontroller CAN entry protection | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [7 empty] | WARNING - The Puritan Bennett™ 5x0 requires special precautions for electromagnetic compatibility and should be installed and started according to the recommendations found in Appendix B, “Specifications.” In particular, the use of nearby mobile and portable communications equipment using radio frequencies, such as mobile telephones or other systems exceeding the levels set in the IEC 60601-1-2 standard, may affect its operation. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [7 empty] | WARNING - The use of any accessory other than those specified, with the exception of the power supplies or cables sold by Covidien, may lead to an increase in electromagnetic emissions or a decrease in the equipment protection against electromagnetic emissions. If the ventilator is used adjacent to such accessories or stacked with such devices, the ventilator’s performance should be monitored to verify normal operation. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD1 |  |  |  |  | The PB560 shall comply with Medical electrical equipment Part 2-55: Particular requirements for basic safety and essential performance of respiratory gas monitor - EN ISO 80601-2-55. |  | 560 WARNING: The ventilator must not use, nor be connected to any anti-static or electrically conductive hoses, tubing, or conduits | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD210 | [6 empty] | WARNING: PC connection is not intended for connection to any devices other than a PC with dedicated compatible Puritan Bennett software package. WARNING: Do not connect the ventilator to any device other than a PC with a dedicated compatible Puritan Bennett software package.
[5 empty] | N |  |  |  | PRD1 | [6 empty] | WARNING: Care should be taken to avoid any potential significant risks of reciprocal interference posed by the ventilator and accessories during specific investigations or treatments.
[5 empty] | N | N |  |  | PRD1 |  |  |  |  | The o2 sensor cable has an overmould cover to protect metal connectirs from ESD (Electrostatic discharge)
[5 empty] | N | N |  |  | PRD1 |  |  |  |  | The upper housing has foil shielding and a conductive gasket to protect against radiated immunity
[5 empty] | N | N |  |  | PRD1a | [6 empty] | Warning: Do not operate the ventilator in a magnetic resonance imaging (MRI) environment. Doing so could cause a ventilator malfunction. Do not operate the ventilator in the presence of active high frequency (HF) surgical equipment. Doing so could cause a ventilator malfunction.
[5 empty] | N | N |  |  | PRD1a | [6 empty] | WARNING: Portable RF communications equipment (including peripherals such as antenna cables and external antennas) should be used no closer than 30 cm (12 inches) to any part of the PB560 ventilator, including cables specified by the manufacturer. Otherwise, degradation of the performance of this equipment could result.
[5 empty] | N | N |  |  | PRD1a | [6 empty] | WARNING: This equipment has been tested and found to comply with the EMC limits for IEC 60601-1-2 (EN 60601-1-2), the EMC Collateral Standard. These limits are designed to provide reasonable protection against harmful interference in a typical medical installation. The equipment generates, uses, and can radiate radio frequency energy and, if not installed and used in accordance with these instructions, may cause harmful interference to other devices in the vicinity or may cause degradation of the performance of this equipment. However, there is no guarantee that interference will not occur in a particular installation.
VH82 | Release of pressurized gas | Gas tanks fall and break at release valve causing pressurized gas release | Use error - tank not restrained | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. • Not suitable for use in the presence of flammable anesthetic mixtures | - |  | 1 | 2 | 1 | 1 | P2 scores based on clinical experience, engineering expertise and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PRD444 | [6 empty] | Manual - WARNING: The ventilator must not be used with flammable anesthetic substances Service Manual - Fire Hazard Warnings • To avoid a fire hazard, keep matches, lighted cigarettes, and all other sources of ignition (e.g., flammable anesthetics and/or heaters) away from the ventilator and oxygen hoses. • In case of fire or a burning smell, immediately disconnect the ventilator from the oxygen supply, facility power, and battery. • Do not use oxygen hoses that are worn, frayed, or contaminated by combustible materials such as grease or oils. (Textiles, oils, and other combustibles are easily ignited and burn with great intensity in air enriched with oxygen.) • The ventilator must not be used with flammable anesthetic substances. • Never expose any batteries to direct flame. | 1 | 2 | 1 | 1
[5 empty] | N | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | PRD444 | [6 empty] | WARNING: To reduce the risk of a fire hazard, keep matches, lighted cigarettes, and all other sources of ignition (such as flammable anesthetics and/or heaters) away from the ventilator and oxygen hoses. | 1 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD229 | [6 empty] | Description of Oxygen function and installation | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD429 | [6 empty] | WARNING: The coupler must not remain connected to the oxygen connector unless it also connected to a leak-proof, external oxygen gas source. When an oxygen supply is not being used with the ventilator, disconnect the oxygen source completely from the ventilator. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD254 | [6 empty] | WARNING: The Puritan Bennett™ 560 Ventilator can be used with an optional oxygen analyzer with minimum and maximum concentration alarms. Always measure the delivered oxygen with a calibrated oxygen analyzer (FiO2 kit) that features a minimum and Installation and Assembly maximum concentration alarm in order to ensure that the prescribed oxygen concentration is delivered to the patient. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD253 | [6 empty] | WARNING: To avoid injury to the patient and/or possible damage to the ventilator: before using the ventilator, use a flow meter (flow regulator) to regulate the oxygen supply to specifications before connecting the ventilator to the oxygen supply. The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD38 |  |  |  |  | The ventilator provides O2 intake port and connector which allows connecting of low pressure oxygen sources at less than 15 lpm at 50 KPa (e.g. Oxygen concentrator). |  |  | 1 | 1 | 1 | 2
|  |  |  | Tissue Damage/Tissue Trauma | N | 2104 | Sudden significant injury to tissue that requires immediate attention to prevent a catastrophic event; Examples: Profuse bleeding on a large vessel (pulmonary artery/vein, renal artery/vein, SMA/IMA); additional unintended tissue loss/resection needed causing additional medical intervention such as a permanent colostomy. Ventilation: Symptomatic; Signs of ARDS requiring Major change in therapy (e.g., ECMO, HFV). | 4 | PRD421 | [6 empty] | WARNING: The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. Inspect the oxygen coupler (Figure 6-16, item 2) before use to ensure it has its black O-ring attached and in good condition. Do not use an oxygen coupler with a missing, damaged, or worn O-ring. Ensure that the oxygen supply pressure to the machine never exceeds 7 psi (50 kPa) or a flow of 15 lpm. Refer to Table A-8. on page A-4 for volume and sensitivity tolerances. | 1 | 2 | 1 | 2
[5 empty] | N | 2104 | Death | 5 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2477 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH83 | Rough surfaces and/or sharp edges (internal or external) | Materials come in contact with patient's, user, manufacturing and service personnel's skin, tissue, or sensitive structures cause abrasion by rubbing or scraping (e.g. friction injury). | User Error - Physical damage from impact or vibration, or transport bag fault | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the Classifications as detailed in IEC 60601-1: • Class II Equipment • Internally Powered Equipment • Type BF Applied Parts • IPX1 with respect to ingress of moisture (IPX2 for PB560 only) • Not suitable for use in the presence of flammable anesthetic mixtures • Not suitable for sterilization • Suitable for continuous operation • Detachable power supply cord |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User injuiry from patient touching display |  | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | PRD8 |  |  |  |  | The PB560 shall comply with Lung Ventilators for Medical Use- Particular Requirements for Basic Safety and Essential Performance Part 2: Home Care Ventilators for Ventilator-Dependent Patient EN ISO 10651-2. |  |  | 1 | 2 | 1 | 1
|  |  | User injuiry from patient touching battery |  | N | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Examples: Bleeding vessels requiring suture repair with blood transfusions; additional/unintended tissue resection required to achieve intended patient outcome. Procedure changed to open procedure to manage bleeding. Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | PRD314 |  |  |  |  | The ventilator shall comply with sections 21.6a (vibration sinusoidal), 21.6b (random vibration) and 21.6c (bump test) of Lung ventilator for medical use - Part 3: Particular requirements for emergency transport ventilators ISO 10651-3. |  |  | 1 | 2 | 1 | 1
|  |  |  | Pain | N | 1994 | discomfort, self-resolving | 1 | PRD364 |  |  |  |  | The ventilator shall be compatible with drop test according to 10.1.3 d article of IEC 60601-1-11 MEDICAL ELECTRICAL EQUIPMENT – PART 1-11: General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment (document IEC/SC62A 62A/624/CD) 2008. |  |  | 1 | 2 | 1 | 1
|  |  | User injuiry from patient touching housing |  | N | 1994 | Slight increase of heart rate, resp. rate. Pain requiring medication via mouth; Pain limits activities of daily living; May include acute pain; TENS Units Symptomatic indicators of minor to moderate pain; Pain requiring non-parenteral medication. | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
|  |  | User injuiry from patient power supply |  | N | 1994 | Severe pain; limits self care; Ultrasound guided injections; intrathecal injections. Major symptoms including increasing vital signs, increasing agitation, analgesics/sedation required, increase ventilator support) | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH84 | Slip Hazard | Inadvertent loss of balance due to fluid spilled on the floor from the ventilator or it's accessories. | Use error - spill of liquid from water trap | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the following product Classifications as detailed in EC 60601-1: • Class II Equipment • Internally Powered Equipment • Type BF Applied Parts • IPX1 with respect to ingress of moisture (IPX2 for PB560 only) • Not suitable for use in the presence of flammable anesthetic mixtures • Not suitable for sterilization • Suitable for continuous operation • Detachable power supply cord |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | PRD8 |  |  |  |  | The PB560 shall comply with Lung Ventilators for Medical Use- Particular Requirements for Basic Safety and Essential Performance Part 2: Home Care Ventilators for Ventilator-Dependent Patient EN ISO 10651-2. |  |  | 1 | 2 | 1 | 1
[5 empty] | N | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Examples: Bleeding vessels requiring suture repair with blood transfusions; additional/unintended tissue resection required to achieve intended patient outcome. Procedure changed to open procedure to manage bleeding. Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | PRD314 |  |  |  |  | The ventilator shall comply with sections 21.6a (vibration sinusoidal), 21.6b (random vibration) and 21.6c (bump test) of Lung ventilator for medical use - Part 3: Particular requirements for emergency transport ventilators ISO 10651-3. | - |  | 1 | 2 | 1 | 1
|  |  |  | Pain | N | 1994 | discomfort, self-resolving | 1 | PRD364 |  |  |  |  | The ventilator shall be compatible with drop test according to 10.1.3 d article of IEC 60601-1-11 MEDICAL ELECTRICAL EQUIPMENT – PART 1-11: General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment (document IEC/SC62A 62A/624/CD) 2008. |  |  | 1 | 2 | 1 | 1
[5 empty] | N | 1994 | Slight increase of heart rate, resp. rate. Pain requiring medication via mouth; Pain limits activities of daily living; May include acute pain; TENS Units Symptomatic indicators of minor to moderate pain; Pain requiring non-parenteral medication. | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N | 1994 | Severe pain; limits self care; Ultrasound guided injections; intrathecal injections. Major symptoms including increasing vital signs, increasing agitation, analgesics/sedation required, increase ventilator support) | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH85 | Sparking | Spark associated with the device or accessory | sparking N/A as it requires an oxygen rich environment to cause a fire which would be a double fault . This double failure is considered under V38
VH86 | Deleted
VH87 | Inspiratory time too short | Device delivers a breath early to patient | User Error - Incorrect set up by Clinician - too sensitive settings | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD132 |  |  |  |  | The system shall increase the inspiratory pressure of at least 0.1 mbar from the PEEP level within 200 ms after a detectable patient effort. After stabilization of the exhalation bias flow as defined in PRD36, some auto cycling could be tolerated according to the following: OK = Possible to trigger an inspiration with increasing effort from sensitivity setting 0P to 5. AC = Ventilator will have high potential to Auto Cycle C= Lung Compliance |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD84 | [6 empty] | WARNING: The inspiration trigger threshold should be carefully modified in order to avoid the risk of false triggering or “autotriggering” of the ventilator. For example, Level 0P, the most sensitive mode, is recommended for pediatric use. However, for an adult, this setting may result in autotriggering. | 2 | 2 | 1 | 1
[5 empty] | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD320 | [5 empty] | Leakage alarms HIGH LEAKAGE Alarm EXH VALVE LEAKAGE ALARM E SENS FAULT OR CIRC LEAK |  | 2 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
[5 empty] | N | 1916 | Death | 5 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) |  | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 2 | 2 | 1 | 1
|  |  |  | Dyspnea |  | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 2 | 1 | 1 | 2
VH88 | Ventilator software upload | Service Engineer unable to upload new revision software | USB memory device hardware fault | Delay of Treatment (prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD395 |  |  |  |  | The service software shall be able to upload ventilator software. |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Labeling Warning in the user manual instructing the user to wait the temperature stabilization of the device before using it after a transport or storage period PRD442 for the PB560-520 [1] |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
|  |  | DFMEA - CPU PCBA: Power Supply Board connector Serial Bus |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
|  |  | DFMEA Power Supply PCBA - USB B (Convertor USB /2UART) |  |  | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
|  |  | SWFMEA PB500 VTS Failure to download software to the ventilator. |  |  | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 1 | 1 | 1 | 2
VH89 | Smoke Hazard | Smoke associated with the device or accessory | DFMEA Power Supply PCBA - ACDC supply O2 SUPPLY - O2 valve oxygen connector - Connecting to the ventilator | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PRD 1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. • Not suitable for use in the presence of flammable anesthetic mixtures |  | Manual - WARNING: The ventilator must not be used with flammable anesthetic substances | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA battery |  | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PRD444 | [6 empty] | WARNING: Never expose any batteries to direct flame. | 1 | 2 | 1 | 1
[5 empty] | N | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | [7 empty] | WARNING: To reduce the risk of a fire hazard, keep matches, lighted cigarettes, and all other sources of ignition (such as flammable anesthetics and/or heaters) away from the ventilator and oxygen hoses. | 1 | 1 | 1 | 1
[5 empty] | N | 2530 | Thermal injury not recognized at the time of surgery. Results in the patient having to undergo urgent surgical intervention due to leakage of a structure/organ. (May need to see peritonitis, aroma, sepsis, etc.) | 4 | PRD439 | [6 empty] | WARNING: Even though the Puritan Bennett™ 560 Ventilator meets current safety standards, the internal Lithium-ion battery of the device exceeds the 100Wh threshold and is therefore considered to be Dangerous Goods (DG) Class 9 – Miscellaneous, when transported in commerce. As such, the Puritan Bennett™ 560 Ventilator and/or the associated Lithium-ion battery are subject to strict transport conditions under the Dangerous Goods Regulation for air transport (IATA: International Air Transport Association), International Maritime Dangerous Goods code for sea and the European Agreement concerning the International Carriage of Dangerous Goods by Road (ADR) for Europe. Private individuals who transport the device are excluded from these regulations although for air transport some requirements apply. For air transport; the Puritan Bennett™ 560 Ventilator is permitted as checked-in or carry-on baggage. Two spare batteries per person may be taken on board as carry-on luggage only, with the prior approval of the airline. This classification and regulatory requirements may vary depending upon the country and mode of transport. Therefore it is recommended that users verify with the carrier / airline as to which measures to take before the voyage. | 1 | 1 | 1 | 2
[5 empty] | N | 2530 | Death | 5 | PRD253 | [6 empty] | WARNING: To avoid injury to the patient and/or possible damage to the ventilator: before using the ventilator, use a flow meter (flow regulator) to regulate the oxygen supply to specifications before connecting the ventilator to the oxygen supply. The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD421 | [6 empty] | WARNING: The hose connecting the ventilator to the oxygen source must be designed exclusively for use with medical-grade oxygen. Under no circumstances should the oxygen hose be modified by the user. In addition, the hose must be installed without the use of lubricants. Inspect the oxygen coupler (Figure 6-16, item 2) before use to ensure it has its black O-ring attached and in good condition. Do not use an oxygen coupler with a missing, damaged, or worn O-ring. Ensure that the oxygen supply pressure to the machine never exceeds 7 psi (50 kPa) or a flow of 15 lpm. Refer to Table A-8. on page A-4 for volume and sensitivity tolerances. | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD429 | [6 empty] | WARNING: The coupler must not remain connected to the oxygen connector unless it also connected to a leak-proof, external oxygen gas source. When an oxygen supply is not being used with the ventilator, disconnect the oxygen source completely from the ventilator. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD214 | [6 empty] | WARNING: Do not attempt to replace the battery yourself. Replacement of lithium batteries or fuel cells by inadequately trained personnel could result in a hazard. Replacement must be completed by qualified service personnel only.
VH90 | Trip Hazard | Device or any of its accessories that occupies space within the user's or caregiver's walkway that may lead to tripping | DFMEA Tubing management - Patient Circuit | Fall | N | 1848 | Minor with no resultant injuries; intervention not indicated | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the following product Classifications as detailed in IEC 60601-1: • Class II Equipment • Internally Powered Equipment • Type BF Applied Parts • IPX1 with respect to ingress of moisture (IPX2 for PB560 only) • Not suitable for use in the presence of flammable anesthetic mixtures • Not suitable for sterilization • Suitable for continuous operation • Detachable power supply cord |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 1848 | Symptomatic; noninvasive intervention indicated | 2 | PRD8 |  |  |  |  | The PB560 shall comply with Lung Ventilators for Medical Use- Particular Requirements for Basic Safety and Essential Performance Part 2: Home Care Ventilators for Ventilator-Dependent Patient EN ISO 10651-2 |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 1848 | Hospitalization indicated indicated for injury related to fall; monitoring | 3 | PRD428 | [6 empty] | WARNING: Do not leave power cables lying on the ground where they may pose a hazard. | 2 | 2 | 1 | 1
[5 empty] | N | 1848 | Urgent surgical treatment indicated due to fall with head injury resulting in a hematoma (See hematoma) needing immediate evacuation to prevent permanent injury. | 4 | PRD19 |  |  |  |  | The ventilator shall be designed such that it can be carried with a strap or carrying case. |  |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD20 |  |  |  |  | The ventilator shall be designed such that it can be placed at the bedside (freestanding on a nightstand). To do so, the dimensions of the ventilator shall be no more than 250 x 350 mm on the horizontal plane.
[5 empty] | N |  |  |  | PRD21 |  |  |  |  | The ventilator shall be designed such that it can be securely mounted on a wheelchair or walker.
[5 empty] | N |  |  |  | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit.
VH91 | Unable to begin ventilation | Ventilation will not begin. Software or device related. E.g. Post, unable to exit startup routine. | Device fails to restart after transport / storage | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD50 PRD53 | [5 empty] | DEVICE FAULT9 alarm - POST RAM Error. RAM Read/Write does DEVICE FAULT10 alarm - POST FLASH Checksum Error. Startup FLASH computed checksum does not match memory setting. DEVICE FAULT11 alarm - POST EEPROM Error. Startup EEPROM does not match memory setting. DEVICE FAULT12 alarm -POST Reference Voltage Error. 5V or 10V reference voltage error. |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use Error - unable to navigate between difference screen/settings |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD395 |  |  |  |  | The service software shall be able to upload ventilator software. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD50 | [6 empty] | Warning: It takes approximately 15 sec to reach essential performance and start ventilation from power on. | 2 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD14 | [6 empty] | WARNING: When the ambient temperature is 20°C, 2 hours are required to both warm the vent from the minimum storage temperature and to cool the vent from maximum storage temperature prior to use | 2 | 2 | 1 | 1
[7 empty] | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | [8 empty] | 2 | 2 | 1 | 1
[7 empty] | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | [8 empty] | 2 | 2 | 1 | 1
[7 empty] | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | 1816 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH92 | Data recorded incorrectly | USB sticks mixed up, wrong information uploaded by user(s). | Use Error | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD401 |  |  |  |  | The clinical software includes a warning popup message to prevent the user from recording the same data file on two different patient profiles and from associating one device to several patient profiles. |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD391 |  |  |  |  | The clinical software only includes the patient ID number, not the patient name. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD210 |  |  |  |  | The clinician software provides capability of retrieving data files through memory devices or directly from the vent through an USB connection. |  |  | 2 | 2 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD334 |  |  |  |  | The clinical software is able to identify each data file coming form the ventilator with its ventilator serial number. |  |  | 2 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD354 | [6 empty] | WARNING: USB connections are not intended for connection to any devices other than the specified USB flash storage (per Section 5.2.1).
VH93 | Use Error - selection of patient type/size, circuit, or information that impacts volume delivery | Breathing circuit and/or settings for breathing circuit chosen inappropriately for size of patient | Use Error - incorrect patient circuit chosen | Dyspnea | Y | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD350 | [5 empty] | Low pressure alarms CHECK EXH VALVE CHECK PROXIMAL LINE1 NO PROXIMAL LINE2 |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use Error - incorrect patient settings input |  | Y | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD179 (560) | [5 empty] | LOW VTE ALARM |  | 2 | 2 | 1 | 1
[5 empty] | Y | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD 181 | [5 empty] | LOW VTI ALARM |  | 2 | 1 | 1 | 2
[5 empty] | Y | 1816 | Death | 5 | PRD178 (560) | [5 empty] | HIGH VTE ALARM |  | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | PRD 180 | [5 empty] | HIGH VTI ALARM |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD244 | [6 empty] | WARNING: To ensure proper performance of the ventilator, use a patient circuit recommended by Covidien in this manual; refer to chapter 6, “Installation and Assembly” and Appendix H, “Parts and Accessories”. The total specified length of the patient circuit tubing as measured from the ventilator outlet to the ventilator inlet is 1.1 meters (3.6 feet) to 2.0 meters (6.6 feet). The tubing must conform to all applicable standards and must be fitted with Ø 22 mm terminals that also conform to all applicable standards. Ensure that both the length and the internal volume of the patient circuit are appropriate for the tidal volume: a corrugated tube of Ø 22 mm for adult patients, and a corrugated tube of Ø 15 mm for pediatric patients with a tidal volume lower than 200 ml. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD246 | [6 empty] | WARNING: Do not start ventilation until you ensure that the device is suitably assembled, that the air inlet filter is properly installed and is not obstructed, and that there is proper clearance all around the unit. Also ensure that the patient circuit is suitably connected to both the ventilator and the patient and that the patient circuit, including all hoses, is not damaged or obstructed. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD230 | [6 empty] | Description of patient circuit | 2 | 1 | 1 | 2
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD231 | [6 empty] | Users Manual - Table F-2. List of Circuits Clinicians Manual - Table H-2. List of Circuits Clinician's Manual - WARNING: For pediatric use, ensure that the patient circuit type fits, and, in all respects, is suitable for use with a child. Use a pediatric circuit for patients that weigh under 53 lb. (23 kg). See Table H-2, List of Circuits on page H-2 for a list of recommended patient circuits. | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD365 |  |  |  |  | When used in conjunction with the DAR exhalation valves from circuits 5093600, 5093500, 5093900, 5094000, 10103098, 10103100, 10103102, The ventilator incompatible with patient circuits with the following characteristics (in addition to standards compliance) : • Length 2.4 meters +/- 10% • Inspiratory resistance at 60 lpm : maximum 2 mb for adults circuits • Inspiratory resistance at 30 lpm : maximum 2 mb for pediatric circuits |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD413 | [6 empty] | Manual - WARNING: Adding accessories to the ventilator breathing circuit, such as a humidifier and water trap(s), may result in a decrease in tidal volume delivered to the patient due to the added compressible volume of the accessory. Always assure that the patient is receiving the appropriate inspired volume when altering the breathing circuit configuration. | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | PRD150 PRD151 |  |  |  |  | Design - Enter Key provides validation of the modification of a setting If the operator does not accept or cancel a change, a time out occurs which will be equivalent to the operator cancellation |  |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD147 |  |  |  |  | The ventilator provides audible feedback to the operator to acknowledge any Key press.
[5 empty] | N |  |  |  | PRD148 |  |  |  |  | The ventilator provides feedback to the operator if they attempt to set a parameter to a value that is not allowed by the ventilator.
[5 empty] | N |  |  |  | PRD149 |  |  |  |  | The ventilator provides a visual distinction between parameters that can be modified by the operator and those that cannot be modified by the operator.
[5 empty] | N |  |  |  | PRD409 | [6 empty] | WARNING: For pediatric use, ensure that the patient circuit type fits, and, in all respects, is suitable for use with a child. Use a pediatric circuit for patients that weigh under 53 lb. (23 kg). See Table H-2, List of Circuits on page H-2 for a list of recommended patient circuits.
[5 empty] | N |  |  |  | PRD411 | [6 empty] | WARNING: The level of inspiratory resistance of the circuit and accessories (bacteria filter, humidifier, HMEs etc.) must be as low as possible. Settings – particularly the PATIENT DISCONNECTION alarm, maximum inspired volume (Max VTI), and minimum inspired volume (Min VTI) settings – must be periodically adjusted according to changes in the patient circuit resistance – especially when filters are replaced.
[5 empty] | N | [8 empty] | Design Locking of the control panel - The Locking Key is a software function that prohibits access to the ventilation and alarm parameter settings and changes to the ventilation mode. |  | NOTE: The ventilator screen must be unlocked before setting and parameters can be changed.
[5 empty] | N | [8 empty] | Ability for the clinician to set the ventilation modes suitable for the patient prior to releasing for homecare use.
[5 empty] | N |  |  |  | PRD414 |  |  |  |  | The ventilator user and clinician manuals shall include a warning instructing the user to check the patient circuit complies to the diameter and length requirements (1 to 2 meters length, 15 or 22 mm diameter). The diameter of 15 mm is recommended for volumes lower than 200 ml. |  | WARNING: To ensure proper performance of the ventilator, use a patient circuit recommended by Covidien in this manual; refer to chapter 6, “Installation and Assembly” and Appendix H, “Parts and Accessories”. The total specified length of the patient circuit tubing as measured from the ventilator outlet to the ventilator inlet is 1.1 meters (3.6 feet) to 2.0 meters (6.6 feet). The tubing must conform to all applicable standards and must be fitted with Ø 22 mm terminals that also conform to all applicable standards. Ensure that both the length and the internal volume of the patient circuit are appropriate for the tidal volume: a corrugated tube of Ø 22 mm for adult patients, and a corrugated tube of Ø 15 mm for pediatric patients with a tidal volume lower than 200 ml.
[5 empty] | N |  |  |  | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD434 | [6 empty] | WARNING: Before dispensing the ventilator to caregivers or the patient for home use, ensure the Locking Key is activated so that critical ventilator settings are not modified.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD272 (560) | [6 empty] | PB560 - Labelling - Ventilator is labelled with 'Refer to instructions manual/booklet' symbol (ISO 7010 - M002)
[5 empty] | N |  |  |  | PRD459 | [6 empty] | WARNING: Before starting ventilation, ensure that the device is properly assembled and that the air inlet, cooling vents, and alarm sound diffusion holes are not obstructed. Ensure also that the patient circuit is of the proper configuration (double or single limb), properly connected to the ventilator, and that the circuit hoses are neither damaged nor compressed and contain no obstructions or foreign bodies.
VH94 | Use Error - Improper use by non-trained personnel | Non trained personnel monitoring and/or setting patient parameters on the ventilator | Misuse of machine: The user has an incorrect knowledge about the device abilities and intended use. | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User misinterruption of ventilator icons |  | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD296 | [6 empty] | Clinician's Manual - Appendix A - Patient/Caregiver Checklist Users Manual - Appendix C - Operational Verification Checklist | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD431 | [6 empty] | WARNING: Be aware this material describes how to respond to ventilator, but does not tell you how to respond to the patient. | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD434 | [6 empty] | WARNING: Before dispensing the ventilator to caregivers or the patient for home use, ensure the Locking Key is activated so that critical ventilator settings are not modified. | 2 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD228 | [6 empty] | WARNING: The ventilator must be used according to its intended use. Refer to section 2.1, “Indications for Use.” | 2 | 2 | 1 | 1
[5 empty] | N |  | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD237 | [6 empty] | WARNING: Before operating the ventilator, read, understand, and strictly follow the information contained in Chapter 1, “Safety Information.” | 2 | 2 | 1 | 1
[5 empty] | N |  | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD462 | [6 empty] | Product Labeling: Prescription by doctor symbol on product label | 2 | 2 | 1 | 1
[5 empty] | N |  | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD272 | [6 empty] | WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
[5 empty] | N | 1816 | Death | 5 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD272 (560) | [6 empty] | PB560 - Labelling - Ventilator is labelled with 'Refer to instructions manual/booklet' symbol (ISO 7010 - M002) | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD293 | [6 empty] | WARNING: To reduce the likelihood of disconnection and to prevent adverse ventilator performance, use only accessories compatible with the ventilator. Compatibility is determined by reviewing the instructions for use of either the ventilator or the accessories”. WARNING: Refer to this manual for equipment compatible with this ventilator. It may be unsafe to interconnect this equipment with other equipment not described in this manual | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD433a | [6 empty] | WARNING: The ventilator system is not intended to be a comprehensive monitoring device and does not activate alarms for all types of conditions. For a detailed understanding of ventilator operations, be sure to thoroughly read this manual before attempting to use the ventilator system. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD433a | [6 empty] | WARNING: The Min PIP alarm setting must be adjusted for the patient, but must also be set high enough to allow the Patient Disconnection alarm to trigger properly. Perform the low pressure test (see section F.1, Low Pressure Test) to ensure that the alarm is properly set. | 2 | 1 | 1 | 2
VH95 | Use Error - Inappropriate Selection of Ventilator settings | Prescribed ventilation or ventilator settings that are not appropriate or are not recommended for the patient condition being treated | Patient condition change, (excessive ventilation or not sufficient ventilation) | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD250 | [6 empty] | Adjustable alarms should not be systematically cancelled; instead, they should be adjusted according to the needs and condition of the patient. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use Error: User accidentally changes settings of the ventilator resulting in inappropriate ventilatory support |  | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD147 |  |  |  |  | The ventilator provides audible feedback to the operator to acknowledge any Key press. |  |  | 2 | 2 | 1 | 1
|  |  | User misinterruption of ventilator icons |  | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD148 |  |  |  |  | The ventilator provides feedback to the operator if they attempt to set a parameter to a value that is not allowed by the ventilator. |  |  | 2 | 1 | 1 | 2
|  |  | User intentionally enters incorrect settings |  | N | 2484 | Death | 5 | PRD149 |  |  |  |  | The ventilator provides a visual distinction between parameters that can be modified by the operator and those that cannot be modified by the operator. |  |  | 2 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD150 PRD151 |  |  |  |  | Design - Enter Key provides validation of the modification of a setting If the operator does not accept or cancel a change, a time out occurs which will be equivalent to the operator cancellation |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD296 | [6 empty] | Clinician's Manual - Appendix A - Patient/Caregiver Checklist Users Manual - Appendix C - Operational Verification Checklist | 2 | 2 | 1 | 1
[5 empty] | Y | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD176 | [5 empty] | HIGH PRESSURE alarm HIGH baseline pressure alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD350 | [5 empty] | Low pressure alarms CHECK EXH VALVE CHECK PROXIMAL LINE1 NO PROXIMAL LINE2 |  | 2 | 1 | 1 | 2
[5 empty] | Y | 1816 | Death | 5 | PRD179 (560) | [5 empty] | LOW VTE ALARM |  | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | PRD 181 | [5 empty] | LOW VTI ALARM |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD178 (560) | [5 empty] | HIGH VTE ALARM |  | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD 180 | [5 empty] | HIGH VTI ALARM |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [5 empty] | Ability for the clinician to set the ventilation modes suitable for the patient prior to releasing for homecare use. |  |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | [5 empty] | Design Locking of the control panel - The Locking Key is a software function that prohibits access to the ventilation and alarm parameter settings and changes to the ventilation mode. |  | NOTE: The ventilator screen must be unlocked before setting and parameters can be changed. | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD434 | [6 empty] | WARNING: Before dispensing the ventilator to caregivers or the patient for home use, ensure the Locking Key is activated so that critical ventilator settings are not modified.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N | [10 empty] | PB560 WARNING: The CPAP mode does not provide a set respiratory rate. Do not use this mode for ventilator dependent patients
[5 empty] | N |  |  |  | PRD419 | [6 empty] | PB560 WARNING: When changing the mode during ventilation, significant transitions of pressure, flow or cycling rate might occur, depending on the difference between the modes. Before setting the new mode, first ensure that the settings between the different modes are compatible. This reduces the risk of discomfort and harm to the patient.
[5 empty] | N |  |  |  | PRD174 | [6 empty] | PB560 WARNING: The setting of the Min PIP alarm must be adjusted for the patient, but must also be set high enough to allow the PATIENT DISCONNECTION alarm to trigger properly. Perform the Low Pressure Test to ensure the Min PIP alarm is properly set.
[5 empty] | N |  |  |  | PRD425 | [6 empty] | PB560 WARNING: Setting any Alarm limits to OFF or extreme high or low values can cause the associated alarms not to activate during ventilation, which reduces its efficacy for monitoring the patient and alerting the clinician to situations that may require intervention.
[5 empty] | N |  |  |  | PRD54 | [6 empty] | PB560 WARNING: A continuous alarm condition will be activated if the ventilator power switch is turned off while ventilation is in progress. When the power switch is turned back on again, the ventilation will resume without having to press the VENTILATION ON/OFF button.
[5 empty] | N |  |  |  | PRD433 | [6 empty] | PB560 WARNING: Before starting ventilation, always verify that all settings are properly set in accordance with the required prescription.
[5 empty] | N |  |  |  | PRD272 (560) | [6 empty] | PB560 - Labelling - Ventilator is labelled with 'Refer to instructions manual/booklet' symbol (ISO 7010 - M002)
[5 empty] | N |  |  |  | PRD65a | [6 empty] | WARNING: In the SIMV mode the use of a double-limb circuit is recommended. The Min VTE setting should remain active in the event that pressure losses are present on the patient circuit downstream from the proximal pressure link. In such cases the "PATIENT DISCONNECTION" alarm would not be systematically activated in case of a disconnection of the circuit
[5 empty] | N |  |  |  | PRD419 | [6 empty] | WARNING: In adult or pediatric use ensure that the adjusted tidal volume is compatible with the needs of the patient.
[5 empty] | N |  |  |  | PRD433 | [6 empty] | WARNING: The ventilator offers a variety of breath delivery options. Throughout the patient’s treatment, the clinician should carefully select the ventilation mode and settings to use for that patient, based on clinical judgment, the condition and needs of the patient, and the benefits, limitations, and characteristics of the breath delivery options. As the patient’s condition changes over time, periodically assess the chosen modes and settings to determine whether those are best for the patient’s current needs.
[5 empty] | Y |  |  |  | PRD174 | [6 empty] | Clinician's Manual: The setting of the LOW PIP alarm must be adjusted for the patient, but must also be set high enough to allow the “PATIENT DISCONNECTION” alarm to trigger properly. Perform the Low Pressure Test (refer to section F.1, “Low Pressure Test” on page F-1) to ensure the LOW PIP alarm is properly set.
VH96 | Use Error - Inappropriate use of the ventilator for transport | Ventilator is used for emergency transport when this is not included in the intended use of this device. | Use Error - inappropriate selection of ventilator for Emergency Use | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | [7 empty] | Manual statement - Target Environments The ventilator is suitable for use in institutional, home, and portable settings. It is not intended for use as an emergency transport ventilator. | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
[5 empty] | N | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[5 empty] | N | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD272 (560) | [6 empty] | PB560 - Labelling - Ventilator is labelled with 'Refer to instructions manual/booklet' symbol (ISO 7010 - M002) | 1 | 1 | 1 | 2
[6 empty] | 1816 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH97 | Use Error - Water trap not emptied | Water trap used in ventilator breathing circuit fills up affecting patient's breathing | DFMEA Water traps - Emptying of water traps | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD412 | [5 empty] | - | WARNING: Always position a humidification device so that it is lower than both the ventilator and the patient. Use water traps, if necessary, to limit water in the patient circuit and precautions should be taken when discarding the fluid in the water trap, discard per local ordinance for proper disposal. If a heated humidifier is used, you should always monitor the temperature of the gas delivered to the patient. Gas delivered from the ventilator that becomes too hot may burn the patient's airway. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD413 | [6 empty] | WARNING: Adding accessories to the ventilator breathing circuit, such as a humidifier and water trap(s), may result in a decrease in tidal volume delivered to the patient due to the added compressible volume of the accessory. Always assure that the patient is receiving the appropriate inspired volume when altering the breathing circuit configuration. | 2 | 2 | 1 | 1
[5 empty] | Y | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 2 | 1 | 1 | 2
[5 empty] | Y | 1816 | Death | 5 | PRD 180 PRD181 PRD188 | [5 empty] | HIGH VTI ALARM LOW VTI Alarm |  | 2 | 1 | 1 | 2
|  |  |  | Pneumothorax | N | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated | 1 | PRD230 | [6 empty] | Description of patient circuit | 2 | 2 | 1 | 1
[5 empty] | N | 2012 | Symptomatic intervention indicated (e.g. tube placement without sclerosis (no pleurodesis), hospitalization observation, electively place chest tube | 2 | PRD234 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. | 2 | 2 | 1 | 1
[5 empty] | N | 2012 | Sclerosis and/or operative intervention indicated; hospitalization indicated (Pleurodesis) Ventilation: Pneumothorax requiring placement chest tube with hospitalization or prolonged hospitalization | 3 | PRD260 | [6 empty] | WARNING: Before opening the packaging for the Patient Circuit, ensure that no damage is evident to the packaging or its contents. Do not use if evidence of damage exists. | 2 | 1 | 1 | 1
|  |  |  | Pneumonia | N | 2011 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated | 1 | PRD295 |  |  |  |  | The pneumatic system design prevents humidity from affecting VTE monitoring performance. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 2011 | Symptomatic; Requires minor increase in ventilatory support, Chest Physical Therapy, antibiotics | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 2011 | Symptomatic; Requires major increased ventilatory support, Antibiotic therapy, and one or more of the following. Chest Physical Therapy and nebulizer around the clock | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD272 (560) | [6 empty] | PB560 - Labelling - Ventilator is labelled with 'Refer to instructions manual/booklet' symbol (ISO 7010 - M002) | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH98 | User Error - Reduced caregiver vigilance | User or caregiver does not maintain or monitor the patient to prescribed settings | User or caregiver does not maintain or monitor the patient to prescribed settings | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD404 | [6 empty] | When an alarm condition is triggered, or there is evidence of a patient-ventilator fault or problem, examine the patient first before examining the ventilator. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User or caregiver does not performed the required checks in advance of connecting a patient to the ventilator |  | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD431 | [6 empty] | WARNING: Be aware this material describes how to respond to ventilator, but does not tell you how to respond to the patient. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD435 | [6 empty] | PB560 - WARNING: Do not perform ventilator alarm tests/maintenance while the patient is connected to the ventilator. Provide the patient with an alternate means of ventilation before conducting these tests. Ensure device is powered off and not in use before routine maintenance is carried out. | 2 | 2 | 1 | 1
[5 empty] | N |  | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | PRD272 (560) | [6 empty] | PB560 - Labelling - Ventilator is labelled with 'Refer to instructions manual/booklet' symbol (ISO 7010 - M002) | 2 | 2 | 1 | 1
[5 empty] | N |  | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | PRD459 | [6 empty] | WARNING: Before starting ventilation, ensure that the device is properly assembled and that the air inlet, cooling vents, and alarm sound diffusion holes are not obstructed. Ensure also that the patient circuit is of the proper configuration (double or single limb), properly connected to the ventilator, and that the circuit hoses are neither damaged nor compressed and contain no obstructions or foreign bodies. | 2 | 2 | 1 | 1
[5 empty] | N |  | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | PRD293 | [6 empty] | WARNING: To reduce the likelihood of disconnection and to prevent adverse ventilator performance, use only accessories compatible with the ventilator. Compatibility is determined by reviewing the instructions for use of either the ventilator or the accessories”. WARNING: Refer to this manual for equipment compatible with this ventilator. It may be unsafe to interconnect this equipment with other equipment not described in this manual | 2 | 1 | 1 | 2
[5 empty] | N | 1816 | Death | 5 | PRD433a | [6 empty] | WARNING: Do not pause, disable, or decrease the volume of the ventilator’s audible alarm if patient safety could be compromised. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH99 | Vent Asynchrony | Ventilator is not able to meet patients work of breathing demands | User Error - Circuit Check is enabled while the patient is being ventilated or connected to the ventilator. | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD459 |  |  |  |  | Circuit leak test |  |  | 1 | 3 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | User Error - ventilator settings inappropriate for patient (e.g. I sense, E sense) |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD460 PRD461 |  |  |  |  | The circuit check menu displays instructions for the steps needed to conduct the test and a confirmation step and a warning to ensure the patient is not connected prior to conducting the test. |  |  | 1 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD162 | [5 empty] | graphics to display patient response |  | 1 | 1 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD237 | [6 empty] | WARNING: Before operating the ventilator, read, understand, and strictly follow the information contained in Chapter 1, “Safety Information.” | 1 | 1 | 1 | 2
[5 empty] | N | 1916 | Death | 5 | PRD413 | [6 empty] | WARNING: Adding accessories to the ventilator breathing circuit, such as a humidifier and water trap(s), may result in a decrease in tidal volume delivered to the patient due to the added compressible volume of the accessory. Always assure that the patient is receiving the appropriate inspired volume when altering the breathing circuit configuration. | 1 | 1 | 1 | 2
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 1 | 3 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD289 | [5 empty] | Leak is displayed |  | 1 | 2 | 1 | 1
[5 empty] | Y | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD320 | [5 empty] | Leakage alarms HIGH LEAKAGE Alarm EXH VALVE LEAKAGE ALARM E SENS FAULT OR CIRC LEAK |  | 1 | 1 | 1 | 1
[5 empty] | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
[6 empty] | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH100 | Ventilator fails to power off | Ventilator does not respond to the shut down sequence | DFMEA Power Supply PCBA Supply switch On/Off switch | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA KEYBOARD MEMBRANE - user interface | No patient involvement | N | 2645 |  | 1 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 2 | 1 | 1
VH101 | Ventilator Inoperative- during use | Ventilator shuts down. | Water Ingress | Hypoventilation | Y | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA Turbine Control PCBA Turbine Power Blower Control Turbine control Battery Connection PCBA - Filtering BLOWER - Turbine Casing/cover Turbine cable to CPU |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD51 |  |  |  |  | Safety net |  |  | 2 | 2 | 1 | 1
|  |  | DFMEA Cables - CPU to Turbine XP Power Supply - ACDC transformer Power Supply - Internal Battery 12V car connection cable - Power to ventilator Misc. Hardware - Cooling fan ACDC Power - cables - ports |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD299 |  |  |  |  | Design - power switch cover to prevent inadvertent power off |  |  | 2 | 2 | 1 | 1
|  |  | DFMEA INSPIRATORY BLOCK Screen Inlet Block Pneumatic connections Overpressure Valve Exhalation Control Valve Exhalation Control Valve assembly and patient circuit to exhalation block Exhalation control valve tube to pressure sensor on the inspiratory port Exhalation block - flow sensor connection |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the following product Classifications as detailed in Clause 5 of IEC 60601-1: • IPX1 with respect to ingress of moisture 560 Only: • IPX2 with respect to ingress of moisture |  |  | 2 | 1 | 1 | 2
|  |  | DFMEA - CPU PCBA: No alarm tone on reset Microcontroller Clock Microcontroller RPD (input pin) Microcontroller Microcontroller 24V MES Microcontroller power supply decoupling Supply board feedback RAM memory Clock Parameters memory Supply +5V +5V Reference +10V Reference Current protection Internal pressure sensor measure Exhalation valve pressure sensor measure (proximal flow pressure sensor) exhalation valve command Turbine connector Exhalation valve connector Exhalation valve connector |  | N | 1916 | Death | 5 | PRD192 | [5 empty] | ALARM FAILURE Alarm |  | 2 | 1 | 1 | 2
|  |  | DFMEA POWER SUPPLY PCBA Supply switch 24 V Management (Step Up/Down Voltage transformer) 5 V supply Micro-controller CPU connector Turbine supply (turbine board supply) current control for battery charge (constant current generator) ACDC supply battery connector contacts | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD167 | [5 empty] | Alarms indicating ventilation failure ignore cancellation attempts |  | 2 | 2 | 1 | 1
|  |  | Blower Control KEYBOARD MEMBRANE - USER INTERFACE O2 SUPPLY - O2 valve oxygen connector - Connecting to the ventilator |  | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD52 |  |  |  |  | Following power up or any CPU reset, the ventilator executes Power On Self Test. |  |  | 2 | 2 | 1 | 1
|  |  | SWFMEA Alarms - System Error-Software Exception Monitors - Incorrect watchdog usage Incorrect battery level monitor |  | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD315 | [5 empty] | INTENTIONAL VENT STOP alarm |  | 2 | 1 | 1 | 2
|  |  | Insufficient mechanical strength or heat resistance |  | N | 2484 | Death | 5 | PRD150 PRD151 |  |  |  |  | Design - Enter Key provides validation of the modification of a setting If the operator does not accept or cancel a change, a time out occurs which will be equivalent to the operator cancellation |  |  | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | N | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD152 |  |  |  |  | The ventilator provides a way to prevent from involuntary ventilation stop by a validation process to acknowledge this command. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD404 | [6 empty] | PB560 WARNING: If a problem with the ventilator is suspected, FIRST CHECK THAT THE PATIENT IS NOT IN DANGER. If necessary, remove the patient from the ventilator and provide an alternative means of ventilation. | 2 | 1 | 1 | 2
[6 empty] | 2477 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH102 | Ventilator Inoperative | Ventilator displays Vent Inop prior to patient use. | DFMEA - CPU PCBA Microcontroller Clock Microcontroller RPD (input pin) Microcontroller microcontroller configuration microcontroller buzzer command chip selsect output extension Microcontroller 24V MES Microcontroller power supply decoupling | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 3 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Supply board feedback RAM memory serial bus Clock Parameters memory Supply +5V +5V Reference +10V Reference Current protection Internal pressure sensor measure |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD51 |  |  |  |  | Safety net |  |  | 2 | 2 | 1 | 1
|  |  | Exhalation valve pressure sensor measure (proximal flow pressure sensor) Exhalation Valve command Turbine connector Exhalation valve connector Exhalation valve connector Microcontroller configuration Microcontroller buzzer command Parameters memory Chip select Output Extension Current protection ACDC supply |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD299 |  |  |  |  | Design - power switch cover to prevent inadvertent power off |  |  | 2 | 1 | 1 | 1
|  |  | DFMEA - Power Supply PCBA - Supply switch 24 V Management (Step Up/Down Voltage transformer) 5 V supply Micro-controller CPU connector Turbine supply (turbine board supply) On/Off switch battery connector contacts |  | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the following product Classifications as detailed in Clause 5 of IEC 60601-1: • IPX1 with respect to ingress of moisture 560 Only: • IPX2 with respect to ingress of moisture |  |  | 2 | 1 | 1 | 2
|  |  | DFMEA - Turbine Control PCBA - Turbine Power - Blower Control - turbine control Battery Connection PCBA - Filtering |  | N |  |  |  | PRD192 | [5 empty] | ALARM FAILURE Alarm
|  |  | Cables - CPU to Turbine BLOWER - Turbine - Casing/cover - Turbine cable to CPU INSPIRATORY BLOCK - Screen - Inlet Block - Pneumatic connections - Overpressure Valve |  | N |  |  |  | PRD167 |  |  |  |  | Alarms indicating ventilation failure ignore cancellation attempts
|  |  | Exhalation Control Valve - Exhalation Control Valve assembly and patient circuit to exhalation block - Exhalation control valve tube to pressure sensor on the inspiratory port Exhalation block - flow sensor connection |  | N |  |  |  | PRD52 |  |  |  |  | Following power up or any CPU reset, the ventilator executes Power On Self Test.
|  |  | Cooling fan Buzzer PCBA - PWM signal process - Buzzer watch function - buzzer amplifier |  | N |  |  |  | PRD315 | [5 empty] | INTENTIONAL VENT STOP alarm
|  |  | POWER SUPPLY internal battery XP Power Supply - ACDC transformer KEYBOARD MEMBRANE - USER INTERFACE |  | N |  |  |  | PRD150 PRD151 |  |  |  |  | Design - Enter Key provides validation of the modification of a setting If the operator does not accept or cancel a change, a time out occurs which will be equivalent to the operator cancellation
|  |  | BLOWER - turbine O2 SUPPLY - O2 valve oxygen connector - Connecting to the ventilator |  | N |  |  |  | PRD152 |  |  |  |  | The ventilator provides a way to prevent from involuntary ventilation stop by a validation process to acknowledge this command.
|  |  | SWFMEA Alarms - System Error-Software Exception Monitors - Incorrect watchdog usage Incorrect battery level monitor |  | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
|  |  | Insufficient mechanical strength or heat resistance |  | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD426 | [6 empty] | PB560 WARNING: If you cannot determine the cause of a problem with your ventilator, contact your equipment supplier. Do not use the ventilator until the problem has been corrected.
[5 empty] | N |  |  |  | PRD426 | [6 empty] | PB560 WARNING: If the ventilator is damaged or its external housing is not correctly closed or it behaves in a way that is not described in this manual (excessive noise, heat emission, unusual odor, alarms not triggered during the start-up procedure), the oxygen and power supplies should be disconnected and use of the device stopped immediately.
VH103 | Ventilator reset/power cycle | The ventilator resets or power cycles during use. | Controls (ON/OFF switch, Ventilation START/STOP, alarm control key, UP/+, Down/-, Menu, Validation) accidentally bumped or activated. | Anxiety | Y | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Machine not properly serviced |  | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD51 |  |  |  |  | Safety net |  |  | 1 | 1 | 1 | 1
|  |  | DFMEA - CPU PCBA - RAM Memory Power Supply PCBA - Micro-controller |  | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD299 |  |  |  |  | Design - power switch cover to prevent inadvertent power off |  |  | 1 | 1 | 1 | 1
|  |  | SWFMEA Alarms - System Error-Software Exception Monitors - Incorrect watchdog usage | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD1 |  |  |  |  | The ventilator shall comply with Medical Electrical Equipment: General Requirements for Safety IEC 60601-1. The ventilator will be constructed to comply with the following product Classifications as detailed in Clause 5 of IEC 60601-1: • IPX1 with respect to ingress of moisture 560 Only: • IPX2 with respect to ingress of moisture |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD167 | [5 empty] | Alarms indicating ventilation failure ignore cancellation attempts |  | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD52 |  |  |  |  | Following power up or any CPU reset, the ventilator executes Power On Self Test. |  |  | 1 | 1 | 1 | 1
[5 empty] | N |  |  |  | PRD315 | [5 empty] | INTENTIONAL VENT STOP alarm
[5 empty] | N |  |  |  | PRD150 PRD151 |  |  |  |  | Design - Enter Key provides validation of the modification of a setting If the operator does not accept or cancel a change, a time out occurs which will be equivalent to the operator cancellation
[5 empty] | N |  |  |  | PRD152 |  |  |  |  | The ventilator provides a way to prevent from involuntary ventilation stop by a validation process to acknowledge this command.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH104 | Battery level indicator is not available | Displayed Battery level indicator is not readable or displayed. | SWFMEA: - Incorrect battery level monitor | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD155 |  |  |  |  | Design - Battery Power indicator flashes when battery is charging |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA Incorrect battery level monitor |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD156 |  |  |  |  | Real time battery life indicator will adjust the indication of battery percent available depending on current ventilator settings and/or setting changes |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD157 |  |  |  |  | Real time battery life indicator will adjust the indication of battery time remaining depending on current ventilator settings and/or setting changes |  | Manual - WARNING: Before using the ventilator’s internal battery, ensure that the battery is fully charged and that the charge holds. Back up ventilators or those in storage should be connected to an AC power source to protect the integrity of the battery. | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD158 |  |  |  |  | The ventilator provides a visual indication of the active ventilator power source (AC, DC or Battery) that is operational in all ventilator states and modes. |  | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD219 |  |  |  |  | When the Ventilator is in Ventilation Stand-by State, the time to fully charge the internal battery from a complete discharge does not exceed 6 hours. |  |  | 1 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD220 |  |  |  |  | When the Ventilator is in Active Ventilation State, the time to fully charge the internal battery from a complete discharge does not exceed 13 hours. | Battery charge level is displayed in time whilst ventilating and in % while on standby |  | 1 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD357 |  |  |  |  | PM schedule - battery charge level information is tested during PM |  |  | 1 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD243 | [6 empty] | Manual - WARNING: Before using the ventilator’s internal battery, ensure that the battery is fully charged and that the charge holds. Back up ventilators or those in storage should be connected to an AC power source to protect the integrity of the battery. | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD423 | [6 empty] | WARNING: The maximum recommended shelf life of the internal battery is two (2) years. Do not use a battery that has been stored for two years prior to its first use. WARNING: Periodic recharging is important to help maximize useful life of the battery. Do not store the internal battery for extended periods, without recharging, as this may reduce the maximum life.
[5 empty] | N |  |  |  | PRD377 |  |  |  |  | The ventilator packaging shall include the storage duration recommended before recharging the internal batteries.
[5 empty] | Y |  |  |  | PRD199 | [5 empty] | Battery alarms EMPTY BATTERY ALARM - Internal battery capacity < 10 min. or 3% CHECK BATTERY CHARGE ALARM - Internal battery charging failure
[5 empty] | Y |  |  |  | PRD309 | [5 empty] | LOW BATTERY alarm | Manual - WARNING: When the “LOW BATTERY“ alarm is triggered, immediately connect the ventilator to an AC power supply to maintain ventilation and recharge the internal battery.
[5 empty] | Y |  |  |  | PRD197 | [5 empty] | AC POWER DISCONNECTION alarm
[5 empty] | N | [10 empty] | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients.
[5 empty] | Y |  |  |  | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH105 | Breathing circuit disconnect | Breathing circuit has a large enough leak to activate circuit disconnect alarm or circuit becomes disconnected from the ventilator or patient | DFMEA Tubing management - Patient Circuit | Hypoventilation | Y | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD350 | [5 empty] | Low pressure alarms CHECK EXH VALVE CHECK PROXIMAL LINE1 NO PROXIMAL LINE2 |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | Y | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD174 | [5 empty] | PATIENT DISCONNECTION alarm |  | 2 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD273 | [5 empty] | All device operational connectors and ports are labeled |  | 2 | 2 | 1 | 1
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD248 | [6 empty] | Manual WARNING: The patient circuit should always be positioned to avoid hindering the patient's movements, to prevent accidental disconnection or leakage, and to minimize the risk of patient strangulation. | 2 | 1 | 1 | 2
[5 empty] | Y | 1916 | Death | 5 | PRD178 (560) PRD179 (560) | [5 empty] | HIGH VTE ALARM LOW VTE Alarm | Manual - WARNING: Minimum and maximum VTE alarm parameters must be properly set to warn in the event of patient disconnection. | 2 | 1 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) | Y | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | PRD 180 PRD181 PRD188 | [5 empty] | HIGH VTI ALARM LOW VTI Alarm |  | 2 | 2 | 1 | 1
[5 empty] | N | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | PRD184 | [5 empty] | NO PROXIMAL LINE2 alarm |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | PRD320 | [5 empty] | Leakage alarms HIGH LEAKAGE Alarm EXH VALVE LEAKAGE ALARM E SENS FAULT OR CIRC LEAK |  | 2 | 1 | 1 | 1
[5 empty] | N | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | PRD287 |  |  |  |  | HIGH LEAK adjustment range |  |  | 2 | 1 | 1 | 2
[5 empty] | N | 2477 | Death | 5 | PRD289 |  |  |  |  | The ventilator displays the inadvertent flow leak in the range |  |  | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD189 (560) | [5 empty] | HIGH RATE alarm (due to autotriggering) |  | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD240 | [6 empty] | Manual - WARNING: After assembling, cleaning, or reassembling the patient circuit, and on a daily basis, inspect the hoses and other components to ensure that there are no cracks or leaks and that all connections are secure. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD306 | [6 empty] | Manual - WARNING: The patient circuit should not be changed during ventilation. | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD230 | [6 empty] | Description of patient circuit | 2 | 1 | 1 | 2
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PRD415 | [6 empty] | Manual - WARNING: When using non-invasive ventilation (NIV), without an exhalation valve, use a vented nose or face mask or a non vented combined with a leak accessory. When using noninvasive ventilation (NIV), with an exhalation valve, use a non vented mask. | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PRD245 | [6 empty] | WARNING: The patient circuit is intended for single use by a single patient and should be changed according to the manufacturer’s recommendations and according to the patient circuit lifetime. | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PRD280 |  |  |  |  | Circuit packaging is ISTA-2 complaint |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [5 empty] | Design - Water traps are screwed in and keyed |  |  | 2 | 1 | 1 | 2
[5 empty] | N | 2328 | See cause of anxiety | 5 | [7 empty] | WARNING: Users must always possess an additional breathing circuit and exhalation valve | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [7 empty] | The circuit check menu displays instructions for the steps needed to conduct the test and a confirmation step and a warning to ensure the patient is not connected prior to conducting the test. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | Design - Hose connections are designed that the fittings are not easily pulled out, optional swivel adapter for patient wye |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD247 | [6 empty] | WARNING: To ensure correct and lasting operation of the ventilator, ensure that its air circulation holes (main inlet or cooling) are never obstructed. Place the device in an area where air can freely circulate around the ventilator and avoid installing it near floating fabrics, such as curtains. Do not place the ventilator in a position where a child, pet or pest can reach it or in any position that might cause it to fall on the patient or someone else. Ensure that the ventilator’s immediate surroundings allow for the proper operational connection of the device without folding, pinching, or damaging any of the required cables or tubes, and that the connection of the patient circuit to the patient provides for a secure, comfortable fit. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD459 |  |  |  |  | Circuit leak test |  |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
[5 empty] | N |  |  |  | PRD411 | [6 empty] | WARNING: The level of inspiratory resistance of the circuit and accessories (bacteria filter, humidifier, and so on) must be as low as possible. Settings – particularly the PATIENT DISCONNECTION alarm, High inspired volume (High VTI), and Low inspired volume (Low VTI) settings – must be periodically adjusted according to changes in the patient circuit resistance – especially when filters are replaced.
VH106 | Displayed data is incorrect | Patient or device (monitored) displayed data displayed value being incorrect is not accurate. May lead to a setting change. | SWFMEA UI - LCD does not update or incorrect data is displayed | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD404 | [6 empty] | When an alarm condition is triggered, or there is evidence of a patient-ventilator fault or problem, examine the patient first before examining the ventilator. | 2 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use Error - Inability to monitor patient data values |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD431 | [6 empty] | WARNING: Be aware this material describes how to respond to ventilator, but does not tell you how to respond to the patient. | 1 | 2 | 1 | 1
|  |  | Incorrect ventilator settings made by clinician due to incorrect pressure, flow and volume displayed |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 1 | 2 | 1 | 1
|  |  | Use Error - Incorrect set-up of patient circuit |  | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 1 | 1 | 1 | 2
[5 empty] | Y | N/A | Death | 5 | [6 empty] | Protect: Depending on the failure, visual and audible alarms appropriate for the condition will alarm. |  | 1 | 1 | 1 | 2
|  |  |  | Dyspnea | N | 1816 | Mild Symptoms, shortness of breath. No treatment required. | 1 | PRD405 | [6 empty] | PB560 WARNING: Ensure that the exhalation block is completely dried after cleaning and prior to use. | 1 | 3 | 1 | 1
[6 empty] | 1816 | Ventilation: Symptomatic; requiring small changes in ventilatory support/oxygen. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1816 | Ventilation: requiring major changes in ventilatory support/oxygen. | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy. | 4 | [8 empty] | 1 | 1 | 1 | 2
VH107 | Inspiratory time too long | Leaks or other anomalies prevents or prolongs transition to exhalation. | Use Error - settings incorrect for patient | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD173 | [5 empty] | HIGH RATE alarm |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA Watchdog Microcontroller BUZZER TEST MES Chip select Output Extension 5V ref failure 10V ref failure 3.3V ref failure Proximal pressure sensor measure Power Supply Board connector |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD110 PRD111 | [5 empty] | low inspiratory tidal volume alarm LOW VTI alarm |  | 1 | 2 | 1 | 1
|  |  | Power Supply PCBA Supply switch 24 V Management (Step Up/Down Voltage transformer) Current control for battery charge (Constant current Generator) 3.3 V supply to USB 5 V supply reference 3 V supply Micro-controller ACDC supply DC external supply (Alimentation DC extern) Cooling fan control (ventilator) CPU connector BATTERY CONNECTOR CONTACTS Cooling fan connector DC supply connector AC/DC supply connector |  | Y | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | PRD172 | [5 empty] | Apnea alarms |  | 1 | 2 | 1 | 1
|  |  | Turbine Control PCBA Blower Temperature |  | N | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | PRD162 | [5 empty] | graphics to display patient response |  | 1 | 1 | 1 | 2
|  |  | Battery Connection PCBA Battery connector Filtering |  |  | 1916 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  | XP Power Supply ACDC transformer | Oxygen saturation, Low (Hypoxemia) |  | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | [8 empty] | 1 | 2 | 1 | 1
|  |  | Power Supply Internal Battery |  |  | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 1 | 2 | 1 | 1
|  |  | Cables Buzzer board to CPU cable Nurse call system / CPU cables and connectors |  |  | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  | ENCLOSURES Lower enclosure Battery cover | Dyspnea |  | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | [8 empty] | 1 | 2 | 1 | 1
|  |  | KEYBOARD MEMBRANE USER INTERFACE |  |  | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | [8 empty] | 1 | 2 | 1 | 1
|  |  | Misc. Hardware Cooling fan |  |  | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | [8 empty] | 1 | 1 | 1 | 1
|  |  | Buzzer PCBA Power supply loss management PWM signal process Buzzer amplifier Buzzer watch function |  |  | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | [8 empty] | 1 | 1 | 1 | 2
|  |  | AC/DC power cables ports | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
|  |  | humidifier connection to the breathing circuit |  |  | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
|  |  | WATER TRAPS connection to the breathing circuit |  |  | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  | 12V car connection cable Power to ventilator |  |  | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 1 | 1 | 1 | 2
VH108 | Battery in use LED indicator is not functioning | The power source indicator light will not illuminate when battery is in use or charging. |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD213 |  |  |  |  | Whenever an AC power supply is available, the ventilator recharges the internal battery |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  | If the patient/user does not know they are running on battery, the vent could shut down in time unexpectantly | DFMEA CPU PCBA keyboard connector KEYBOARD MEMBRANE USER INTERFACE |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD217 |  |  |  |  | The ventilator is able to estimate battery ageing. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD219 |  |  |  |  | When the Ventilator is in Ventilation Stand-by State, the time to fully charge the internal battery from a complete discharge does not exceed 6 hours. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD220 |  |  |  |  | When the Ventilator is in Active Ventilation State, the time to fully charge the internal battery from a complete discharge does not exceed 13 hours. | Battery charge level is displayed in time whilst ventilating and in % while on standby |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD157 |  |  |  |  | Real time battery life indicator will adjust the indication of battery time remaining depending on current ventilator settings and/or setting changes |  | Manual - WARNING: Before using the ventilator’s internal battery, ensure that the battery is fully charged and that the charge holds. Back up ventilators or those in storage should be connected to an AC power source to protect the integrity of the battery. | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | Y | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD199 | [5 empty] | Battery alarms EMPTY BATTERY ALARM - Internal battery capacity < 10 min. or 3% CHECK BATTERY CHARGE ALARM - Internal battery charging failure |  | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD243 | [6 empty] | Manual - WARNING: Before using the ventilator’s internal battery, ensure that the battery is fully charged and that the charge holds. Back up ventilators or those in storage should be connected to an AC power source to protect the integrity of the battery. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD423 | [6 empty] | WARNING: The maximum recommended shelf life of the internal battery is two (2) years. Do not use a battery that has been stored for two years prior to its first use. | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD377 |  |  |  |  | The ventilator packaging includes the storage duration recommended before recharging the internal batteries. |  |  | 2 | 1 | 1 | 2
[5 empty] | Y |  |  |  | PRD309 | [5 empty] | LOW BATTERY alarm | Manual - WARNING: When the “LOW BATTERY“ alarm is triggered, immediately connect the ventilator to an AC power supply to maintain ventilation and recharge the internal battery.
[5 empty] | N | [10 empty] | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients.
[5 empty] | Y |  |  |  | PRD197 | [5 empty] | AC POWER DISCONNECTION alarm
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH109 | AC in use LED indicator is not functioning | The power source indicator light will not illuminate when AC power is in use. |  | No consequences or impact to patient | Y | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD197 | [5 empty] | AC POWER DISCONNECTION alarm |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | DFMEA CPU PCBA keyboard connector KEYBOARD MEMBRANE USER INTERFACE | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD158 |  |  |  |  | The ventilator provides a visual indication of the active ventilator power source (AC, DC or Battery) that is operational in all ventilator states and modes. |  | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD299 |  |  |  |  | Design - power switch cover to prevent inadvertent power off |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD241 PRD242 PRD243 | [6 empty] | WARNING: The power supply to which the ventilator is connected (both AC and DC) must comply with all applicable standards and provide electrical power corresponding to the voltage characteristics inscribed on the rear of the ventilator to ensure correct operation. Refer also to the electrical specifications found in Appendix B, “Specifications”. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD227 |  |  |  |  | Design - When operating on external power and with charged internal battery, the ventilator continues to ventilate within its specified performance limits when external power voltage is reduced to any value between its nominal rated voltage and zero volts (brown out). |  |  | 2 | 1 | 1 | 2
[5 empty] | Y | N/A | Death | 5 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD194 PRD195 | [5 empty] | DEVICE FAULT 3 alarm DEVICE FAULT 4 alarm POWER FAULT alarm
[5 empty] | N |  |  |  | PRD258 | [6 empty] | Ensure that the AC and DC power cables are in perfect condition and not compressed. The device should not be turned on if the AC or DC power cable is damaged.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH110 | DC in use LED indicator is not functioning | The power source indicator light will not illuminate when DC power is in use. | DFMEA LED fail | Delay of Treatment (Prolonged procedure) | Y | N/A | Minor delay, no patient impact | 1 | PRD198 | [5 empty] | DC POWER DISCONNECTION alarm |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | SWFMEA Monitors - No main power loss detection |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [7 empty] | Manual - WARNING: To ensure that ventilation continues uninterrupted ensure alternative power sources are available (AC power source, extra batteries, or an auxiliary DC car adapter). Be prepared for the possibility of power failure by having an alternative means of ventilation ready for use—particularly for ventilator-dependent patients. | 2 | 2 | 1 | 1
|  |  | DFMEA CPU PCBA keyboard connector POWER SUPPLY PCBA DC external connector KEYBOARD MEMBRANE USER INTERFACE |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD299 |  |  |  |  | Design - power switch cover to prevent inadvertent power off |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD223 |  |  |  |  | The device is designed to be protected in case power connections are reversed at DC input. |  |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD241 PRD242 PRD243 | [6 empty] | WARNING: The power supply to which the ventilator is connected (both AC and DC) must comply with all applicable standards and provide electrical power corresponding to the voltage characteristics inscribed on the rear of the ventilator to ensure correct operation. Refer also to the electrical specifications found in Appendix B, “Specifications”. | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PRD227 |  |  |  |  | Design - When operating on external power and with charged internal battery, the ventilator continues to ventilate within its specified performance limits when external power voltage is reduced to any value between its nominal rated voltage and zero volts (brown out). |  |  | 2 | 2 | 1 | 1
[5 empty] | Y | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PRD53 PRD54 | [5 empty] | Security shut down alarm will sound on total power loss |  | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PRD194 PRD195 | [5 empty] | DEVICE FAULT 3 alarm DEVICE FAULT 4 alarm POWER FAULT alarm |  | 2 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | PRD351 |  |  |  |  | DC power cable retainer |  |  | 2 | 1 | 1 | 2
[5 empty] | N |  |  |  | PRD217 |  |  |  |  | battery level indicator
[5 empty] | N |  |  |  | PRD258 | [6 empty] | Ensure that the AC and DC power cables are in perfect condition and not compressed. The device should not be turned on if the AC or DC power cable is damaged.
[5 empty] | Y |  |  |  | PRD309 | [5 empty] | LOW BATTERY alarm | Manual - WARNING: When the “LOW BATTERY“ alarm is triggered, immediately connect the ventilator to an AC power supply to maintain ventilation and recharge the internal battery.
[5 empty] | N |  |  |  | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended.
[5 empty] | N |  |  |  | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem.
VH111 | Impaired Capnography function | Device displayed or sensor measured ETCO2 value data is not accurate | N/A for the PB500 series ventilators - no Capnography capability
VH112 | Alternative ventilation delivered | Alternate ventilation delivered in response to ventilator fault, apnea, and/or procedure errors. | Loss of patient drive | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [6 empty] | graphics to display patient response |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | Loss of oxygen |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD430 | [6 empty] | WARNING: This ventilator must be used only under the responsibility and on the prescription of a doctor. | 2 | 2 | 1 | 1
[5 empty] | Y | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [6 empty] | Apnea alarms |  | 2 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PRD433 (560) | [6 empty] | PB560 - WARNING: A ventilator-dependent patient should always be monitored by trained and competent medical personnel. Ensure that the patient’s caregiver is able and prepared to take suitable action in the event the ventilator identifies an alarmed condition or experiences a problem. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD435 | [6 empty] | Manual - Warnings - While the ventilator is in use, an alternative means of ventilation should always be available in the event of a ventilator problem. This is particularly true for ventilator-dependent patients. Supplementary observation, appropriate for the patient's condition, is also recommended. | 2 | 1 | 1 | 2
POWERPACK | [5 empty] | The Puritan Bennett™ Power Pack provides an external battery which connects to the Puritan Bennett™ 560 and 520 Ventilators. It is an external DC power source. Hazards listed below do not include the hazards associated with ventilation but have been selected for issues with the functionality of the PowerPack. Harms and severities will be low as failure of the PowerPack results in switching to the internal ventilator battery.
VH14 | Battery fails to charge | Battery fails to charge (internal or external) |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [7 empty] | WARNING: When recharging the power pack batteries, the power supply to which the external charger is connected must comply with all applicable standards and provide electrical power corresponding to the AC voltage characteristics stated on the external charger label. | 3 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD56 | [5 empty] | DC POWER DISCONNECTION alarm |  | 3 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PRD198 | [5 empty] | Charge indicator push button |  | 3 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PPKPRD28 | [5 empty] | Battery charge indicator LED | There is one LED on the external battery charger. *Orange Start of charge *Yellow Medium level of charge *Green Charge is complete | 3 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PPKPRD29 | [5 empty] | Battery capacity remaining LED |  | 3 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PPKPRD30 | [5 empty] | Battery capacity low LED | When the battery is running low, the farthest-left LED on the battery charge indicator turns yellow/orange. | 3 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PPKPRD31 | [5 empty] | AC mains connected LED |  | 3 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PPKPRD97 | [6 empty] | Each time the charger is plugged into the power pack, there is a start up auto-test. | 3 | 1 | 1 | 2
[5 empty] | N | 2484 | Death | 5 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 3 | 1 | 1 | 2
[5 empty] | N | [10 empty] | Safety checks
VH15 | Battery level indication is inaccurate | Displayed Battery level indicator displays incorrect charge status of battery |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD28 | [5 empty] | Charge indicator push button |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD29 | [5 empty] | Battery charge indicator LED | There is one LED on the external battery charger. *Orange Start of charge *Yellow Medium level of charge *Green Charge is complete | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PPKPRD30 | [5 empty] | Battery capacity remaining LED |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PPKPRD31 | [5 empty] | Battery capacity low LED | When the battery is running low, the farthest-left LED on the battery charge indicator turns yellow/orange. | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PPKPRD97 | [5 empty] | AC mains connected LED |  | 2 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PPKPRD42 |  |  |  |  | The PB560 battery fully recharges in less than 6 hours. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PPKPRD46 |  |  |  |  | The Power Pack turns OFF automatically during the following conditions: • Batteries are depleted • Shutoff due to temperature out of range |  | When the power pack battery is fully depleted, the ventilator automatically switches to the ventilator’s internal battery and an audible and visible DC POWER DISCONNECTION alarm sounds on the ventilator. | 2 | 2 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | PPKPRD102 PPKPRD103 |  |  |  |  | Design - On charging the installed batteries priority is to the fixed battery followed by the removable battery When discharging, the first battery to be discharged is removable battery followed by the fixed battery |  |  | 2 | 1 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH17 | Battery is unable to be removed | Battery is stuck or jammed in place. |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD16 |  |  |  |  | The Power Pack provides a means of latching and unlatching the access panel to a removable battery. |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD17 |  |  |  |  | The power pack provides a place for a removable battery and a fixed battery. |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PPKPRD18 |  |  |  |  | The power pack casing provides a door, so that the removable battery can be removed and replaced (hot-swappable). |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - Include an opening so that the battery can be removed from the casing and “hot-swapped”. |  |  | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PPKPRD114 | [6 empty] | Description of the installation & removal procedure for the removable battery | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | PPKPRD118 | [6 empty] | The Power Pack removable battery door is labelled indicating the correct orientation of battery installation
[5 empty] | N | [8 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails
VH18 | Biohazard Exposure | During the use or handling of the device or system (e.g. device surface, circuit, humidifiers) user or patient is exposed to pathogens. | Using same power pack between patients in a long term or skilled nursing facility. | Infection | N | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | PPKPRD27 | [6 empty] | WARNING: Regularly clean the power pack carrying bag according to manufacturer’s recommendations. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | Improper cleaning/maintenance |  | N | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth, Incision/drainage see abscess | 2 | PPKPRD62 | [6 empty] | Clean all external panels and surfaces before and after each patient use and as often as necessary to keep the battery pack clean. Clean the battery pack weekly, whenever it is soiled or dirty, before any maintenance operation, and before storing. | 2 | 2 | 1 | 1
[5 empty] | N | 1930 | An infection that requires more advanced medical intervention including hospital admission for treatment/observation; IV antibiotics; IV antifungals, IV antivirals; non-emergent scheduled operative intervention (See pneumonia, immunosuppression) | 3 | [6 empty] | Safety check |  | 2 | 1 | 1 | 1
|  |  |  | Cross Patient Exposure To Body Fluids |  | 1745 | No medical treatment required if in contact with in-tack skin | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1745 | prophalacric treatment required for exposure to blood or body fluid | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1745 | prophalacric treatment required for exposure to blood or body fluid | 3 | [8 empty] | 2 | 1 | 1 | 1
VH23 | Burning smell or odor | Device or device accessory exhibits a burning smell or odor | DFMEA PowerPack PCBA Removable Battery damage | Anxiety | N | 2328 | Mild symptoms; intervention Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PPKPRD72 | [6 empty] | WARNING: Never expose batteries to direct flame. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | DFMEA Internal Battery temperature sensor breakdown voltage sensor breakdown Internal Battery short circuit |  | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | PPKPRD59 | [6 empty] | WARNING: Ensure that the AC and DC power cables are in perfect condition and not compressed. The device should not be turned on if the AC or DC power cable is damaged. | 2 | 2 | 1 | 1
|  |  | DFMEA PowerPack Charger Charger failure Short circuit on DC output Short circuit on AC mains input |  | N | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | PPKPRD14 PPKPRD15 PPKPRD55 |  |  |  |  | Operational conditions Transport and storage conditions |  | WARNING: To ensure correct and lasting operation of the device, install and operate the equipment in the environmental conditions recommended in this manual. | 2 | 1 | 1 | 1
|  |  | DFMEA Charger comes loose from mounts DC inlet/outlet port damage DC cable is damaged |  | N | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [6 empty] | Safety check |  | 2 | 1 | 1 | 2
|  |  |  | Chemical (gas or liquid)exposure , Hazardous | N | 2570-Chemical Exposure | Minor Inflammation, no treatment needed | 1 | PPKPRD46 |  |  |  |  | The Power Pack turns OFF automatically during the following conditions: • Batteries are depleted • Shutoff due to temperature out of range |  | When the power pack battery is fully depleted, the ventilator automatically switches to the ventilator’s internal battery and an audible and visible DC POWER DISCONNECTION alarm sounds on the ventilator. | 2 | 2 | 1 | 1
[5 empty] | N | 2570-Chemical Exposure | Temporary injury or impairment that requires minor professional medical intervention | 2 | PPKPRD32 | [5 empty] | Temperature LED |  | 2 | 2 | 1 | 1
[6 empty] | 2570-Chemical Exposure | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | [8 empty] | 2 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 3 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 3 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 5 | [8 empty] | 2 | 1 | 1 | 2
VH24 | Cable Entanglement | Device or device accessory that becomes entangled around patient or user | Cable wraps around neck, legs, arms - patient or user harm | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PPKPRD44 |  |  |  |  | The Power Pack is supplied with a 24V DC cable capable of connecting the unit to the PB520/560 ventilator when the unit is mounted in the carrier bag configuration A second DC cable is available capable of connecting Power Pack to the ventilator, with a length of 80 cm +/- 5cm. |  |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | Power cord wraps around neck, legs, arms - patient harm |  | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. | 2 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 1 | 1 | 1
|  |  | Power cord wraps around neck, legs, arms - ventilator disconnect |  |  | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2104 | Sudden significant injury to tissue that requires immediate attention to prevent a catastrophic event | 4 | [8 empty] | 1 | 1 | 1 | 2
VH26 | Chemical toxicity | Exposure of chemicals, or cleaning agents to patient, user, or environment | Use error - inappropriate cleaning chemicals | Chemical exposure, Hazardous | N | 2570 | Minor Inflammation, no treatment needed | 1 | PPKPRD27 | [6 empty] | WARNING: Regularly clean the power pack carrying bag according to manufacturer’s recommendations. | 2 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | External battery leakage |  | N | 2570 | Temporary injury or impairment that requires minor professional medical intervention | 2 | PPKPRD62 | [6 empty] | Clean all external panels and surfaces before and after each patient use and as often as necessary to keep the battery pack clean. Clean the battery pack weekly, whenever it is soiled or dirty, before any maintenance operation, and before storing. | 2 | 1 | 1 | 1
|  |  | User Error - dropping of the battery or PowerPack |  |  | 2570 | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | [8 empty] | 2 | 1 | 1 | 1
|  |  |  | Tissue Damage/Tissue Trauma |  | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | [8 empty] | 2 | 1 | 1 | 1
VH28 | Contamination | Introduction of debris, dust, or contaminants to patient, user, or environment. | Dual or carrying bag dirty | Infection | N | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | PPKPRD27 | [6 empty] | WARNING: Regularly clean the power pack carrying bag according to manufacturer’s recommendations. | 2 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | PM schedule not adhered to |  | N | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth or inhalation, Incision/drainage see abscess | 2 | PPKPRD26 PPKPRD66 |  |  |  |  | The exterior of the Power Pack is cleanable with a cloth or sponge lightly moistened with a bactericide or germicide solution |  | WARNING: Use all cleaning solutions and products with caution, using only those solutions listed in this manual. Read and follow the instructions associated with the cleaning solutions used to clean the device. | 2 | 1 | 1 | 1
[5 empty] | N | 1930 | Symptomatic; IV antibiotics required with increased ventilatory support. An infection that requires more advanced medical intervention | 3 | PPKPRD62 | [6 empty] | Clean all external panels and surfaces before and after each patient use and as often as necessary to keep the battery pack clean. Clean the battery pack weekly, whenever it is soiled or dirty, before any maintenance operation, and before storing. | 2 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [6 empty] | Safety check |  | 2 | 1 | 1 | 1
VH29 | Damage to internal system components | Internal device system component damage due to oxidation, water ingress, servicing, installation. | Use Error - water ingress | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD54 | [6 empty] | WARNING: Do not operate in direct sunlight, near heat sources, outdoors, or near installations where liquid may pose a risk without first providing adequate protection for the device. | 3 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | Use Error - PowerPack is dropped |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD70 | [6 empty] | WARNING: Never immerse the device in any liquid or allow liquid to enter any device opening. Never place a container of liquid on top of the device. Immediately wipe away any liquid on the surface of the device. To avoid damage to the batteries or electrical components, prevent fluids from entering the device. | 3 | 1 | 1 | 1
|  |  | Use Error - polarity of battery is reversed |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [5 empty] | Design - PowerPack assembly is designed to have a plastic assembly which would protect the battery packs and the electronics |  |  | 3 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PRD421 | [6 empty] | PB560 WARNING: Ensure that the only gas supplied to the ventilator through the dedicated oxygen supply connector is medical-grade oxygen. | 3 | 1 | 1 | 1
VH30 | Device/Device Accessories Become Hot | Any part of device or accessories, where the user handles the device, is hot due to environmental exposure or excessive heating of internal components (e.g. batteries, electrical components). |  | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PPKPRD32 | [5 empty] | Temperature LED |  | 3 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PPKPRD54 | [6 empty] | WARNING: Do not operate in direct sunlight, near heat sources, outdoors, or near installations where liquid may pose a risk without first providing adequate protection for the device. | 3 | 1 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD46 |  |  |  |  | The Power Pack turns OFF automatically during the following conditions: • Batteries are depleted • Shutoff due to temperature out of range |  | When the power pack battery is fully depleted, the ventilator automatically switches to the ventilator’s internal battery and an audible and visible DC POWER DISCONNECTION alarm sounds on the ventilator. | 3 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | Design- the battery has a thermal cut off and temperature sensing. The battery will turn off if it overheats. |  |  | 3 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [5 empty] | Circulation holes in the powerpack housing to allow air to circulate |  |  | 3 | 1 | 1 | 1
VH33 | Electrical shock | The passage of electric current through the body (e.g. exposed electrical components, transmission of electrical current through fluids including bodily). |  | Electrical shock | N | 2554 | Asymptomatic, intervention not indicated. Slight tingling sensation; if burn see Harm of Burn | 1 | [6 empty] | Safety check |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | 2554 | Non-urgent medical intervention indicated. Slight shock felt; not painful but disturbing. Average individual can let go. However, strong involuntary reactions to shocks in this range may lead to injuries .5mA ; if burn see Harm of Burn | 2 | PPKPRD112 |  |  |  |  | The AC mains input is fuse protected |  |  | 1 | 1 | 1 | 1
[5 empty] | N | 2554 | Medical intervention indicated (planned cardioversion)Painful shock, begin to lose muscular control. Commonly referred to as the freezing current or "let-go" range. 6mA to 16mA.; if burn see Harm of Burn | 3 | PPKPRD117 | [6 empty] | Description of the replacement procedure for the AC mains fuse | 1 | 1 | 1 | 1
[5 empty] | N | 2554 | Life-threatening consequences; hemodynamic compromise; urgent intervention indicated; Extreme pain, respiratory arrest, severe muscular contractions. Individual cannot let go. Death is possible. 17mA-99mA; if burn see Harm of Burn | 4 | PPKPRD110 |  |  |  |  | The power pack requires the use of a tool to open the enclosure and access the internal circuitry of the device. |  |  | 1 | 1 | 1 | 2
[5 empty] | N | 2554 | Death +G1422 | 5 | PPKPRD111 |  |  |  |  | The enclosure protects against connection of the internal or removable battery’s terminals with any power pack circuitry when the batteries are installed in an incorrectly |  |  | 1 | 1 | 1 | 2
|  |  |  | Burn, Thermal |  | 2530 | Thermal injury not resulting in need for treatment. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | [8 empty] | 1 | 1 | 1 | 1
VH34 | Environmental specifications exceeded | Ventilator used or stored outside of the environmental specifications as described in the user manual | User Error - environmental temperature or relative humidity exceeded | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD54 | [6 empty] | WARNING: Do not operate in direct sunlight, near heat sources, outdoors, or near installations where liquid may pose a risk without first providing adequate protection for the device. | 1 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD14 PPKPRD15 PPKPRD55 |  |  |  |  | Operational conditions Transport and storage conditions |  | WARNING: To ensure correct and lasting operation of the device, install and operate the equipment in the environmental conditions recommended in this manual. | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PPKPRD69 | [6 empty] | WARNING: Handle the power pack with care during and after use, particularly when ambient temperatures are high. Some power pack surfaces may become hot, even if safety specifications are not exceeded. | 1 | 1 | 1 | 1
[5 empty] | N |  |  |  | PPKPRD46 |  |  |  |  | The Power Pack turns OFF automatically during the following conditions: • Batteries are depleted • Shutoff due to temperature out of range |  | When the power pack battery is fully depleted, the ventilator automatically switches to the ventilator’s internal battery and an audible and visible DC POWER DISCONNECTION alarm sounds on the ventilator.
[5 empty] | N |  |  |  | PPKPRD46 |  |  |  |  | The Power Pack turns OFF automatically during the following conditions: • Batteries are depleted • Shutoff due to temperature out of range |  | When the power pack battery is fully depleted, the ventilator automatically switches to the ventilator’s internal battery and an audible and visible DC POWER DISCONNECTION alarm sounds on the ventilator.
[5 empty] | N | [8 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails
VH38 | Fire Hazard | Fire associated with the device or accessory | DFMEA PowerPack PCBA Removable Battery damage | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PPKPRD72 | [6 empty] | WARNING: Never expose batteries to direct flame. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | DFMEA Internal Battery temperature sensor breakdown voltage sensor breakdown Internal Battery short circuit |  | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PPKPRD59 | [6 empty] | WARNING: Ensure that the AC and DC power cables are in perfect condition and not compressed. The device should not be turned on if the AC or DC power cable is damaged. | 2 | 2 | 1 | 1
|  |  | DFMEA PowerPack Charger Charger failure Short circuit on DC output Short circuit on AC mains input |  | N | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | PPKPRD14 PPKPRD15 PPKPRD55 |  |  |  |  | Operational conditions Transport and storage conditions |  | WARNING: To ensure correct and lasting operation of the device, install and operate the equipment in the environmental conditions recommended in this manual. | 2 | 1 | 1 | 1
|  |  | DFMEA Charger comes loose from mounts DC inlet/outlet port damage DC cable is damaged |  | N | 2530 | Thermal injury not recognized at the time of surgery. Results in the patient having to undergo urgent surgical intervention due to leakage of a structure/organ. (May need to see peritonitis, aroma, sepsis, etc.) | 4 | [6 empty] | Safety check |  | 2 | 1 | 1 | 2
[5 empty] | N | 2530 | Death | 5 | PPKPRD46 |  |  |  |  | The Power Pack turns OFF automatically during the following conditions: • Batteries are depleted • Shutoff due to temperature out of range |  | When the power pack battery is fully depleted, the ventilator automatically switches to the ventilator’s internal battery and an audible and visible DC POWER DISCONNECTION alarm sounds on the ventilator. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD32 | [5 empty] | Temperature LED |  | 2 | 1 | 1 | 1
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [5 empty] | Design- the battery has a thermal cut off and temperature sensing. The battery will turn off if it overheats. |  |  | 2 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH39 | Hazardous materials | Device or device materials when discarded may cause effects to patient, user or environment. |  | Chemical exposure, Hazardous | N | 2570 | Minor Inflammation, no treatment needed | 1 | PPKPRD74 PPKPRD75 | [6 empty] | WARNING: The battery and the power pack should be considered as waste electrical and electronic equipment (WEEE). Dispose of battery and power pack in accordance with local requirements and regulations. | 3 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  | Improper device disposal |  |  | N | 2570 | Temporary injury or impairment that requires minor professional medical intervention | 2 | PPKPRD83 | [6 empty] | The Power pack is labeled to indicate that the device should be considered as waste electrical and electronic equipment. | 3 | 1 | 1 | 1
[5 empty] | N | 2570 | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | [5 empty] | Design - sealed battery case |  |  | 3 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 3 | 1 | 1 | 1
VH50 | Inadequate surge protection | Electrical circuit fails during use due to a power surge. | Damaged power cord used | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD112 |  |  |  |  | AC mains fuse protection |  |  | 2 | 2 | 1 | 1 | P2 scores based on clinical expertise, engineering input, complaint data, and environment of use. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 2 | 1 | 1
VH54 | Loose Internal mechanical interface or connection | Any internal parts, connectors, or components that become loose or disconnected | Use Error - PowerPack is dropped | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD65 | [6 empty] | WARNING: Never use a power pack or any components or accessories that appear to be damaged. If any signs of damage are evident, contact your equipment supplier or Covidien. | 2 | 2 | 1 | 1 | P2 scores based on clinical expertise, engineering input, complaint data, and environment of use. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD67 | [6 empty] | WARNING: To minimize the risk of damage, you must use the carrying bag to transport the power pack. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [6 empty] | Safety check |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PPKPRD12 PPKPRD13 |  |  |  |  | Lung ventilator for medical use - Part 3: Particular requirements for emergency transport ventilators ISO 10651-3. bump, vibration and drop tests |  |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - Be attachable to the dual bag |  |  | 2 | 1 | 1 | 2
[5 empty] | N | [8 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails
VH55 | Loss of external mechanical interface or connection | Any external parts, connectors, or components that become disconnected | User Error - cables not connected correctly | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD65 | [6 empty] | WARNING: Never use a power pack or any components or accessories that appear to be damaged. If any signs of damage are evident, contact your equipment supplier or Covidien. | 2 | 2 | 1 | 1 | P2 scores based on clinical expertise, engineering input, complaint data, and environment of use. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | User Error - Physical damage from impact or vibration |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD67 | [6 empty] | WARNING: To minimize the risk of damage, you must use the carrying bag to transport the power pack. | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PPKPRD12 PPKPRD13 |  |  |  |  | Lung ventilator for medical use - Part 3: Particular requirements for emergency transport ventilators ISO 10651-3. bump, vibration and drop tests |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [6 empty] | Safety check |  | 2 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - Be attachable to the dual bag |  |  | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PPKPRD25 |  |  |  |  | The battery carrying bag is designed such that the Power Pack display is visible and all connection points are accessible. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | [9 empty] | DC DISCONNECTION ALARM
[5 empty] | N | [8 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails
VH69 | Inadequate ventilator maintenance | Machine not properly serviced |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD110 |  |  |  |  | The power pack shall require use of a tool to open the enclosure and access the internal circuitry of the device. |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical experience, engineering expertise and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD113 | [6 empty] | Service requirements | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 2 | 1 | 1
VH72 | No Hazard | No hazard occurred during reported event or identified after event analysis | DFMEA PowerPack PCBA J11 DC output voltage connector open circuit +5V regulator breakdown fails in open circuit: RL1 relay, FL1 filter, Charging /discharging switches T2, T4, T14 switch Microcontroller breakdown +3V regulator breakdown J9, T8, T9, T14 open circuit F1 fuse breakdown PowerPack Internal Battery - voltage sensor breakdown Removable Battery - open circuit PowerPack Charger - connector port to vent DC port is damaged | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [5 empty] | none |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical experience, engineering expertise and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
DFMEA PowerPack Keyboard Key to turn on LED battery level indicator cannot turn on LED battery level indicator cannot turn off PowerPack carry bag Strap failure PowerPack Housing Removable battery door not locked in place or lock mechanism broken damage to the battery slot /housing Damage to the pins / PCBA Internal damage/short circuit on Charger board or Main PCBA - Fan mount screws come loose inside unit AC inlet/outlet sockets fall out of housing - DC inlet/outlet port damage DC cable is damaged Battery Charger Cable - AC Power Cord or connector damaged, wires broken
VH76 | Loose External mechanical interface or connection | Any external parts, connectors, or components that become that become loose. | User Error - Physical damage from impact or vibration | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD67 | [6 empty] | WARNING: To minimize the risk of damage, you must use the carrying bag to transport the power pack. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix |  | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [6 empty] | Safety check |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - Be attachable to the dual bag |  |  | 2 | 2 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PPKPRD34 |  |  |  |  | The DC connector is protected against accidental disconnection. |  |  | 2 | 1 | 1 | 1
[5 empty] | N |  |  |  | PPKPRD100 |  |  |  |  | Power Pack provides a means of protecting any connected AC cables against accidental disconnection |  |  | 2 | 1 | 1 | 1
VH76 | Loose External mechanical interface or connection | Any external parts, connectors, or components that become that become loose. | User Error - Physical damage from impact or vibration | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD67 | [6 empty] | WARNING: To minimize the risk of damage, you must use the carrying bag to transport the power pack. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix |  | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [6 empty] | Safety check |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - Be attachable to the dual bag |  |  | 2 | 2 | 1 | 1
|  |  |  | No consequences or impact to patient | N | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | PPKPRD34 |  |  |  |  | The DC connector is protected against accidental disconnection. |  |  | 2 | 1 | 1 | 1
[5 empty] | N |  |  |  | PPKPRD100 |  |  |  |  | Power Pack provides a means of protecting any connected AC cables against accidental disconnection
VH78 | Ventilator and/or accessories tips, moves or falls | The ventilator or ventilator assembly’s moves, falls or tips on Patient, user, bystander. Impact on environment also possible. | DFMEA PowerPack carry bag Strap failure | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - Be attachable to the dual bag |  |  | 3 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | Wheelchair attachment fails |  |  | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. | 2 | [8 empty] | 3 | 2 | 1 | 1
|  |  | User drops power pack |  |  | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue | 3 | [8 empty] | 3 | 2 | 1 | 1
|  |  |  | Anxiety | N | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | PPKPRD67 | [6 empty] | WARNING: To minimize the risk of damage, you must use the carrying bag to transport the power pack. | 3 | 2 | 1 | 1
[5 empty] | N | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 3 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 3 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 3 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 3 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 3 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 3 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 3 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 3 | 1 | 1 | 2
VH81 | EMC Susceptibility | Interference of the ventilators ability to provide ventilation due to electromagnetic energy. |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD116 | [6 empty] | Description of the electromagnetic emissions, immunity and recommended separation distances in relation to operation of the device in proximity to electromagnetic fields WARNING: Portable and mobile RF communications equipment can affect the performance of the ventilator and power pack. Install and use this device according to the information contained in this manual. Interference may occur in the vicinity of equipment marked with the following symbol: The ventilator and power pack should not be used adjacent to or stacked with other equipment, except as specified in this manual. If adjacent or stacked use is necessary, the ventilator should be observed to verify normal operation in the configurations in which it will be used. | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  | Note this is looking at the effect on the ventilator if the PowerPack fails due to EMC |  |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [7 empty] | WARNING - The Puritan Bennett™ 5x0 requires special precautions for electromagnetic compatibility and should be installed and started according to the recommendations found in Appendix B, “Specifications.” In particular, the use of nearby mobile and portable communications equipment using radio frequencies, such as mobile telephones or other systems exceeding the levels set in the IEC 60601-1-2 standard, may affect its operation. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [7 empty] | WARNING - The use of any accessory other than those specified, with the exception of the power supplies or cables sold by Covidien, may lead to an increase in electromagnetic emissions or a decrease in the equipment protection against electromagnetic emissions. If the ventilator is used adjacent to such accessories or stacked with such devices, the ventilator’s performance should be monitored to verify normal operation. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PRD1 |  |  |  |  | The powerpack AC pass through cable has a ferrite attached to protect against radiated emissions. |  |  | 1 | 1 | 1 | 2
VH83 | Rough surfaces and/or sharp edges (internal or external) | Materials come in contact with patients, user, manufacturing and service personnel's skin, tissue, or sensitive structures cause abrasion by rubbing or scraping (e.g. friction injury). | User Error - Physical damage from impact or vibration, or transport bag fault | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PPKPRD67 | [6 empty] | WARNING: To minimize the risk of damage, you must use the carrying bag to transport the power pack. | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[6 empty] | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Pain |  | 1994 | discomfort, self-resolving | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1994 | Slight increase of heart rate, resp. rate. Pain requiring medication via mouth; Pain limits activities of daily living; May include acute pain; TENS Units Symptomatic indicators of minor to moderate pain; Pain requiring non-parenteral medication. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1994 | Severe pain; limits self care; Ultrasound guided injections; intrathecal injections. Major symptoms including increasing vital signs, increasing agitation, analgesics/sedation required, increase ventilator support) | 3 | [8 empty] | 1 | 2 | 1 | 1
VH88 | Software update not available | Service Engineer unable to upload new revision software | DFMEA PowerPack PCBA Connector breakdown | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [5 empty] | none |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack be unavailable, the ventilator will have the internal battery available .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 1 | 1 | 1 | 2
VH89 | Smoke Hazard | Smoke associated with the device or accessory | DFMEA PowerPack PCBA Removable Battery damage | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PPKPRD72 | [6 empty] | WARNING: Never expose batteries to direct flame. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | DFMEA Internal Battery temperature sensor breakdown voltage sensor breakdown Internal Battery short circuit |  | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PPKPRD59 | [6 empty] | WARNING: Ensure that the AC and DC power cables are in perfect condition and not compressed. The device should not be turned on if the AC or DC power cable is damaged. | 2 | 2 | 1 | 1
|  |  | DFMEA PowerPack Charger Charger failure Short circuit on DC output Short circuit on AC mains input |  | N | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | PPKPRD14 PPKPRD15 PPKPRD55 |  |  |  |  | Operational conditions Transport and storage conditions |  | WARNING: To ensure correct and lasting operation of the device, install and operate the equipment in the environmental conditions recommended in this manual. | 2 | 2 | 1 | 1
|  |  | DFMEA Charger comes loose from mounts DC inlet/outlet port damage DC cable is damaged |  | N | 2530 | Thermal injury not recognized at the time of surgery. Results in the patient having to undergo urgent surgical intervention due to leakage of a structure/organ. (May need to see peritonitis, aroma, sepsis, etc.) | 4 | [6 empty] | Safety check |  | 2 | 1 | 1 | 2
[5 empty] | N | 2530 | Death | 5 | PPKPRD46 |  |  |  |  | The Power Pack turns OFF automatically during the following conditions: • Batteries are depleted • Shutoff due to temperature out of range |  | When the power pack battery is fully depleted, the ventilator automatically switches to the ventilator’s internal battery and an audible and visible DC POWER DISCONNECTION alarm sounds on the ventilator. | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD32 | [5 empty] | Temperature LED |  | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH90 | Trip Hazard | Device or any of its accessories that occupies space within the user's or caregiver's walkway that may lead to tripping | Cables | Fall | N | 1848 | Minor with no resultant injuries; intervention not indicated | 1 | PPKPRD44 |  |  |  |  | The Power Pack is supplied with a 24V DC cable capable of connecting the unit to the PB520/560 ventilator when the unit is mounted in the carrier bag configuration A second DC cable is available capable of connecting Power Pack to the ventilator, with a length of 80 cm +/- 5cm. |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | Misuse of machine: The user has an incorrect knowledge about the device abilities and intended use. |  |  | 1848 | Symptomatic; noninvasive intervention indicated | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1848 | Hospitalization indicated for injury related to fall; monitoring | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1848 | Urgent surgical treatment indicated due to fall with head injury resulting in a hematoma (See hematoma) needing immediate evacuation to prevent permanent injury. | 4 | [8 empty] | 2 | 1 | 1 | 2
VH94 | Use Error - Improper use by non-trained personnel | Non trained personnel monitoring and/or setting patient parameters on the ventilator |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PPKPRD58 | [6 empty] | WARNING: Other than replacing the battery, do not attempt to open, repair or otherwise service the power pack. Doing so might endanger the patient, damage the battery, and/or void the warranty. Only personnel authorized and qualified by Covidien should repair, open, or service the power pack. | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PPKPRD82 | [6 empty] | Power pack outer carton Label contains a symbol reminding the operator to read the user’s manual. Read, understand, and follow these instructions before using the power pack. In addition, to use the power pack correctly and efficiently and to help prevent incidents, review and adhere to all warnings and cautions contained in the Clinician’s or User’s Manual specific to the ventilator in use. | 2 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Dyspnea |  | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | 1816 | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
VH104 | Battery level indicator is not available | Displayed Battery level indicator is not readable or displayed. |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PPKPRD28 | [5 empty] | Charge indicator push button |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD29 | [5 empty] | Battery charge indicator LED | There is one LED on the external battery charger. *Orange Start of charge *Yellow Medium level of charge *Green Charge is complete | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PPKPRD30 | [5 empty] | Battery capacity remaining LED |  | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PPKPRD31 | [5 empty] | Battery capacity low LED | When the battery is running low, the farthest-left LED on the battery charge indicator turns yellow/orange. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PPKPRD97 | [5 empty] | AC mains connected LED |  | 1 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - The battery display must be visible through the carrying bag window |  |  | 1 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PPKPRD25 |  |  |  |  | The battery carrying bag is designed such that the Power Pack display is visible and all connection points are accessible. |  |  | 1 | 1 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 1 | 1 | 1 | 2
PB500 SERIES PORTABLE VENTILATOR CART 10000646 | [5 empty] | Hazards listed below have been selected for issues with the functionality of the cart. This is an optional accessory and not all ventilators are associated with a cart.
VH38 | Fire Hazard | Fire associated with the device or accessory | Cylinder falls from cart | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | [7 empty] | WARNING: To ensure stability, when the Puritan Bennett™ 560 Ventilator is mounted on a cart, the weight of the oxygen bottle should not exceed 14 kg (30 lbs.). | 1 | 2 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
[6 empty] | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2530 | Thermal injury not recognized at the time of surgery. Results in the patient having to undergo urgent surgical intervention due to leakage of a structure/organ. (May need to see peritonitis, aroma, sepsis, etc.) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2530 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH47 | Improper Ergonomic Approach | Service tech or user does not use proper technique in lifting heavy ventilator component. | Use Error - cart is lifted inappropriately | Fall | N | 1848 | Minor with no resultant injuries; intervention not indicated | 1 | [7 empty] | WARNING: To ensure stability, when the Puritan Bennett™ 560 Ventilator is mounted on a cart, the weight of the oxygen bottle should not exceed 14 kg (30 lbs.). | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
[6 empty] | 1848 | Symptomatic; noninvasive intervention indicated | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1848 | Hospitalization indicated for injury related to fall; monitoring | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1848 | Urgent surgical treatment indicated due to fall with head injury resulting in a hematoma (See hematoma) needing immediate evacuation to prevent permanent injury. | 4 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Fatigue |  | 1849 | Minor fatigue that comes and goes | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1849 | Symptomatic; fatigue lasting longer than a few hours or days unrelieved by rest. | 2 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Repetitive Strain/Stress injury |  | 2689 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2689 | Altered function resulting in need for oral medication | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2689 | Severely altered function; hospitalization indicated; elective non-urgent operative intervention indicated | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2689 | Permanent damage to function of extremity; permanent disability | 4 | [8 empty] | 1 | 1 | 1 | 2
VH55 | Loss of external mechanical interface or connection | Any external parts, connectors, or components that become disconnected | User Error - cart not assembled correctly | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD446 | [6 empty] | WARNING: To ensure stability, when the Puritan Bennett™ 560 Ventilator is mounted on a cart, the weight of the oxygen bottle should not exceed 14 kg (30 lbs.). | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
|  |  | User Error - Physical damage from impact or vibration |  |  | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
|  |  | User Error - User does not attach ventilator to cart securely |  |  | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Respiratory failure |  | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH76 | Loose External mechanical interface or connection | Any external parts, connectors, or components that become that become loose. | User Error - cart not assembled correctly | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD446 |  |  |  |  | 4 mounting studs prevent movement of the ventilator on the cart | The cart meets the rough handling, stability, and suspension systems tests per UL 60601-1 with the ventilator attached. | WARNING: To ensure stability, when the Puritan Bennett™ 560 Ventilator is mounted on a cart, the weight of the oxygen bottle should not exceed 14 kg (30 lbs.). | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
|  |  | User Error - Physical damage from impact or vibration |  |  | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
|  |  | User Error - User does not attach ventilator to cart securely |  |  | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
VH78 | Ventilator and/or accessories tips, moves or falls | The ventilator or ventilator assembly’s moves, falls or tips on Patient, user, bystander. Impact on environment also possible. | Cart wheel falls off, uneven floor or roll over something that causes imbalance | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PRD451 |  |  |  |  | 4 mounting studs prevent movement of the ventilator on the cart | The cart meets the rough handling, stability, and suspension systems tests per UL 60601-1 with the ventilator attached. |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
[5 empty] | N | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | PRD451 |  |  |  |  | brakes on 2 wheels | The cart meets the rough handling, stability, and suspension systems tests per UL 60601-1 with the ventilator attached. |  | 1 | 2 | 1 | 1
[5 empty] | N | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Examples: Bleeding vessels requiring suture repair with blood transfusions; additional/unintended tissue resection required to achieve intended patient outcome. Procedure changed to open procedure to manage bleeding. Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | PRD451 |  |  |  |  | Design - low center of gravity | The cart meets the rough handling, stability, and suspension systems tests per UL 60601-1 with the ventilator attached. |  | 1 | 2 | 1 | 1
[5 empty] | N |  |  |  | PRD451 |  |  |  |  | Design - 5 wheels to enhance stability | The cart meets the rough handling, stability, and suspension systems tests per UL 60601-1 with the ventilator attached.
[5 empty] | N |  |  |  | PRD446 | [6 empty] | WARNING: To ensure stability, when the Puritan Bennett™ 560 Ventilator is mounted on a cart, the weight of the oxygen bottle should not exceed 14 kg (30 lbs.).
VH80 | Pressure line (including prox flow tubing) is kinked or obstructed | Impact on ventilatory support due to changes in pressure, flow and/or volume | Use Error- pressure line kinked during set up or during patient movement | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD451 |  |  |  |  | Circuit arm mount bar Humidifier mount |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD446 | [6 empty] | WARNING: To ensure stability, when the Puritan Bennett™ 560 Ventilator is mounted on a cart, the weight of the oxygen bottle should not exceed 14 kg (30 lbs.). | 1 | 1 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 1916 | Death | 5 | [8 empty] | 1 | 2 | 1 | 2
|  |  |  | Pneumothorax |  | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated Asymptomatic-clinical or diagnostic finding. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2012 | Small to Moderate Pneumothorax, electively place chest tube or observation | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2012 | Large Pneumothorax, electively place chest tube | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2012 | Tension Pneumothorax; Emergent Chest tube placement, intubation. | 4 | [8 empty] | 1 | 2 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) |  | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2477 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH82 | Release of pressurized gas | Gas tanks fall and break at release valve causing pressurized gas release |  | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | PRD451 |  |  |  |  | Cylinder holder |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
[5 empty] | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | PRD446 | [6 empty] | WARNING: To ensure stability, when the Puritan Bennett™ 560 Ventilator is mounted on a cart, the weight of the oxygen bottle should not exceed 14 kg (30 lbs.). | 1 | 2 | 1 | 1
[6 empty] | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Tissue Damage/Tissue Trauma |  | 2104 | Sudden significant injury to tissue that requires immediate attention to prevent a catastrophic event; Examples: Profuse bleeding on a large vessel (pulmonary artery/vein, renal artery/vein, SMA/IMA); additional unintended tissue loss/resection needed causing additional medical intervention such as a permanent colostomy. Ventilation: Symptomatic; Signs of ARDS requiring Major change in therapy (e.g., ECMO, HFV). | 4 | [8 empty] | 1 | 2 | 1 | 2
[6 empty] | 2104 | Death | 5 | [8 empty] | 1 | 2 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) |  | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2477 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH83 | Rough surfaces and/or sharp edges (internal or external) | Materials come in contact with patients, user, manufacturing and service personnel's skin, tissue, or sensitive structures cause abrasion by rubbing or scraping (e.g. friction injury). | User Error - Physical damage from impact or vibration, or transport bag fault | Tissue Damage/Tissue Trauma |  | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | [5 empty] | none |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
[6 empty] | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Examples: Bleeding vessels requiring suture repair with blood transfusions; additional/unintended tissue resection required to achieve intended patient outcome. Procedure changed to open procedure to manage bleeding. Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Pain |  | 1994 | discomfort, self-resolving | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1994 | Slight increase of heart rate, resp. rate. Pain requiring medication via mouth; Pain limits activities of daily living; May include acute pain; TENS Units Symptomatic indicators of minor to moderate pain; Pain requiring non-parenteral medication. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1994 | Severe pain; limits self care; Ultrasound guided injections; intrathecal injections. Major symptoms including increasing vital signs, increasing agitation, analgesics/sedation required, increase ventilator support) | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH90 | Trip Hazard | Device or any of its accessories that occupies space within the user's or caregiver's walkway that may lead to tripping | DFMEA Tubing management - Patient Circuit | Fall | N | 1848 | Minor with no resultant injuries; intervention not indicated | 1 | PRD451 |  |  |  |  | Circuit arm mount bar Humidifier mount |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The portable ventilator cart offers a mobility solution for daily activities and outings. The benefits of having a mobile cart for home care ventilation outweighs this potential risk.
[6 empty] | 1848 | Symptomatic; noninvasive intervention indicated | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1848 | Hospitalization indicated for injury related to fall; monitoring | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1848 | Urgent surgical treatment indicated due to fall with head injury resulting in a hematoma (See hematoma) needing immediate evacuation to prevent permanent injury. | 4 | [8 empty] | 1 | 1 | 1 | 2
CARRYING BAG POWERPACK 10000609 PB 540 CARRYING BAG 3829000 PB 540 DUAL BAG 2967200 | [5 empty] | Hazards listed below have been selected for issues with the functionality of the bags. The ventilator is supplied with the carrying bag. The Dual bag is an optional accessory which allows the user to move the ventilator on a wheelchair or carry as a back pack.
VH20 | Breathing circuit pinched or kinked | Patient Breathing circuit has a restriction causing decreased pressure, flow, or volume. | Use error - Patient circuit occluded, pinched or kinked | Hypoventilation | N | 1916 | Asymptomatic; PaCO2 or ETCO2 elevated. | 1 | PRD303 |  |  |  |  | The transport bag (dual bag) is designed in a way that the ventilator display is visible and all accessory connection points are accessible. |  |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The carrying bags offered with the ventilator offers mobility solutions for daily activities and outings. The benefits of having dedicated carry bags for home care ventilation and wheelchair mounting outweighs this potential risk.
[5 empty] | N | 1916 | Symptomatic; PaCO2 or ETCO2 elevated. ↓ in Minute Ventilation, requires minor ventilator adjustment | 2 | PRD255 | [6 empty] | WARNING: To minimize the risk of damage, you must use the Dual Bag to transport the Puritan Bennett™ 560 Ventilator. | 1 | 1 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated. Severe ↓ in Minute Ventilation, requires major ventilator adjustment | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 1916 | Symptomatic; PaCO2 or ETCO2 severely elevated/loss of respiratory drive. Severe ↓ in Minute Ventilation, requires Major change of therapy (e.g., ECMO, CPR, Intubation) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 1916 | Death | 5 | [8 empty] | 1 | 2 | 1 | 2
|  |  |  | Pneumothorax |  | 2012 | Asymptomatic; clinical or diagnostic observations only; intervention not indicated Asymptomatic-clinical or diagnostic finding. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2012 | Small to Moderate Pneumothorax, electively place chest tube or observation | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2012 | Large Pneumothorax, electively place chest tube | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2012 | Tension Pneumothorax; Emergent Chest tube placement, intubation. | 4 | [8 empty] | 1 | 2 | 1 | 2
|  |  |  | Oxygen saturation, Low (Hypoxemia) |  | 2477 | Episodic shortness of breath; resolved with rest. Asymptomatic: self resolving. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2477 | Symptomatic and/or ≤92: Minor increase or incremental increases in oxygen levels resolves | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤88% : Major increase or incremental increases in oxygen levels and increase ventilatory support. | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2477 | Symptomatic ≤85% prolonged or ≤ 80% sudden and continuous dropping : Major increase or incremental increases in oxygen levels and Major increase ventilatory support. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2477 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH28 | Contamination | Introduction of debris, dust, or contaminants to patient, user, or environment. | Dual or carrying bag dirty | Infection | N | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | PRD303 |  |  |  |  | The transport bag (dual bag) is designed in a way that the ventilator display is visible and all accessory connection points are accessible. |  |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The carrying bags offered with the ventilator offers mobility solutions for daily activities and outings. The benefits of having dedicated carry bags for home care ventilation and wheelchair mounting outweighs this potential risk.
[5 empty] | N | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth or inhalation, Incision/drainage see abscess | 2 | PRD360 | [6 empty] | Regularly clean the ventilator’s Dual Bag according to manufacturer’s recommendations. | 1 | 1 | 1 | 1
[6 empty] | 1930 | Symptomatic; IV antibiotics required with increased ventilatory support. An infection that requires more advanced medical intervention | 3 | [8 empty] | 1 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 1 | 1 | 1
VH55 | Loss of external mechanical interface or connection | Any external parts, connectors, or components that become disconnected | User Error - Physical damage from impact or vibration, or transport bag fault | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The carrying bags offered with the ventilator offers mobility solutions for daily activities and outings. The benefits of having dedicated carry bags for home care ventilation and wheelchair mounting outweighs this potential risk.
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 2 | 1 | 1
VH71 | Display field(s) are not visible | Patient or device (monitored) displayed data is not visible to user/clinician | Use Error - Ventilator inserted into the carrying bag incorrectly | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | PRD303 |  |  |  |  | The transport bag (dual bag) is designed in a way that the ventilator display is visible and all accessory connection points are accessible. |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The carrying bags offered with the ventilator offers mobility solutions for daily activities and outings. The benefits of having dedicated carry bags for home care ventilation and wheelchair mounting outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PRD384 |  |  |  |  | The transport bag (dual bag) shall allow insertion of the ventilator in one direction only. |  |  | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 3 | 1 | 2
|  |  |  | Respiratory failure |  | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 1 | 3 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 1 | 3 | 1 | 2
VH76 | Loose External mechanical interface or connection | Any external parts, connectors, or components that become that become loose. | User Error - Physical damage from impact or vibration, or transport bag fault | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The carrying bags offered with the ventilator offers mobility solutions for daily activities and outings. The benefits of having dedicated carry bags for home care ventilation and wheelchair mounting outweighs this potential risk.
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 2 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 2 | 2 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 1 | 1 | 1
VH78 | Ventilator and/or accessories tips, moves or falls | The ventilator or ventilator assembly’s moves, falls or tips on Patient, user, bystander. Impact on environment also possible. | Dual or carrying bag strap fails | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - Be attachable to the dual bag |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[6 empty] | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH83 | Rough surfaces and/or sharp edges (internal or external) | Materials come in contact with patients, user, manufacturing and service personnel's skin, tissue, or sensitive structures cause abrasion by rubbing or scraping (e.g. friction injury). | User Error - Physical damage from impact or vibration, or transport bag fault | Tissue Damage/Tissue Trauma |  | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | [8 empty] | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[6 empty] | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Pain |  | 1994 | discomfort, self-resolving | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1994 | Slight increase of heart rate, resp. rate. Pain requiring medication via mouth; Pain limits activities of daily living; May include acute pain; TENS Units Symptomatic indicators of minor to moderate pain; Pain requiring non-parenteral medication. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1994 | Severe pain; limits self care; Ultrasound guided injections; intrathecal injections. Major symptoms including increasing vital signs, increasing agitation, analgesics/sedation required, increase ventilator support) | 3 | [8 empty] | 1 | 2 | 1 | 1
VH90 | Trip Hazard | Device or any of its accessories that occupies space within the user's or caregiver's walkway that may lead to tripping | DFMEA Tubing management - Patient Circuit | Fall | N | 1848 | Minor with no resultant injuries; intervention not indicated | 1 | PRD451 |  |  |  |  | Circuit arm mount bar Humidifier mount |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[6 empty] | 1848 | Symptomatic; noninvasive intervention indicated | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1848 | Hospitalization indicated for injury related to fall; monitoring | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1848 | Urgent surgical treatment indicated due to fall with head injury resulting in a hematoma (See hematoma) needing immediate evacuation to prevent permanent injury. | 4 | [8 empty] | 1 | 1 | 1 | 2
PB540 PowerPack 4096677 | [5 empty] | The Puritan Bennett™ Power Pack provides an external battery which connects to the Puritan Bennett™ 540 Ventilator. It is an external DC power source. Hazards listed below do not include the hazards associated with ventilation but have been selected for issues with the functionality of the PowerPack. Harms and severities will be low as failure of the PowerPack results in switching to the internal ventilator battery.
VH14 | Battery fails to charge | Battery fails to charge (internal or external) | PCBA malfunctions/damaged | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | N/A |  |  |  |  | Battery charge indicator LED |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Battery non functional |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | N/A | [6 empty] | The battery charger's LED light will be YELLOW while the battery is charging. When the battery is fully charged, the LED light will switch to GREEN. This transition may be accompanied by a brief flashing of the GREEN LED. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | N/A | [6 empty] | Note: To maintain a fully charged battery, the charger may remain connected until ready for use. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | N/A | [5 empty] | DC POWER DISCONNECTION alarm |  | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | N/A | [6 empty] | WARNING Ensure that the ventilator’s internal battery is fully charged before connecting the ventilator to an external DC power source. Powering the ventilator using an external 12– 30 VDC power source (via the DC power cable) does not enable charging of its internal battery. | 1 | 1 | 1 | 2
|  |  |  | Respiratory failure |  | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH17 | Battery is unable to be removed | Battery is stuck or jammed in place. | Zipper on bag is stuck | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | N/A | [6 empty] | Direction: Carefully disconnect the wires from the battery terminals by grasping the plastic wire terminal and pulling; do not pull directly on the wire. | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Connection wires to fuse not releasing |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | N/A | [6 empty] | Caution Do not reverse the battery polarity when connecting the battery kit to the battery. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | N/A |  |  |  |  | Design - zipper and flap holds the battery in place |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | N/A | [6 empty] | Description of the installation & removal procedure for the removable battery | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | N/A |  |  |  |  | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 1 | 1 | 2
[5 empty] | N |  |  |  | N/A | [6 empty] | Manual gives a table indicating the power available - 'A fully charged external battery provides ventilator power for the following minimum durations (- 10%) based on ventilator settings, environmental conditions, and physiological characteristics of the patient: '
VH18 | Biohazard Exposure | During the use or handling of the device or system (e.g. device surface, circuit, humidifiers) user or patient is exposed to pathogens. | Using same power pack between patients in a long term or skilled nursing facility. | Infection |  | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | None | [7 empty] | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Improper cleaning/maintenance |  |  | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth, Incision/drainage see abscess | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1930 | An infection that requires more advanced medical intervention including hospital admission for treatment/observation; IV antibiotics; IV antifungals, IV antivirals; non-emergent scheduled operative intervention (See pneumonia, immunosuppression) | 3 | [8 empty] | 1 | 1 | 1 | 1
|  |  |  | Cross Patient Exposure To Body Fluids |  | 1745 | No medical treatment required if in contact with in-tack skin | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1745 | prophalacric treatment required for exposure to blood or body fluid | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1745 | prophalacric treatment required for exposure to blood or body fluid | 3 | [8 empty] | 1 | 1 | 1 | 1
VH23 | Burning smell or odor | Device or device accessory exhibits a burning smell or odor | Use Error - PowerPack is dropped | Anxiety | N | 2328 | Mild symptoms; intervention Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | N/A | [6 empty] | WARNING: Never expose batteries to direct flame. | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Chemical (gas or liquid)exposure , Hazardous |  | 2570-Chemical Exposure | Minor Inflammation, no treatment needed | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2570-Chemical Exposure | Temporary injury or impairment that requires minor professional medical intervention | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2570-Chemical Exposure | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | [8 empty] | 1 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 3 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 3 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 5 | [8 empty] | 1 | 1 | 1 | 2
VH24 | Cable Entanglement | Device or device accessory that becomes entangled around patient or user | Cable wraps around neck, legs, arms - patient or user harm | Tissue Damage/Tissue Trauma | N | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | N/A |  |  |  |  | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Power cord wraps around neck, legs, arms - patient harm |  |  | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. | 2 | [8 empty] | 1 | 1 | 1 | 1
|  |  | Power cord wraps around neck, legs, arms - ventilator disconnect |  |  | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2104 | Sudden significant injury to tissue that requires immediate attention to prevent a catastrophic event | 4 | [8 empty] | 1 | 1 | 1 | 2
VH26 | Chemical toxicity | Exposure of chemicals, or cleaning agents to patient, user, or environment | Use error - inappropriate cleaning chemicals | Chemical exposure, Hazardous |  | 2570 | Minor Inflammation, no treatment needed | 1 | [5 empty] | none |  |  | 2 | 1 | 1 | 1 | P2 scores are based on clinical expert opinion, engineering inputs, and complaint data | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | External battery leakage |  |  | 2570 | Temporary injury or impairment that requires minor professional medical intervention | 2 | [8 empty] | 2 | 1 | 1 | 1
|  |  | User Error - dropping of the battery or PowerPack |  |  | 2570 | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | [8 empty] | 2 | 1 | 1 | 1
|  |  |  | Tissue Damage/Tissue Trauma |  | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Ventilation: Symptomatic; Early Radiologic findings of Ventilator induced lung injury requiring minor ventilator adjustments. | 2 | [8 empty] | 2 | 1 | 1 | 1
[6 empty] | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue Ventilation: Symptomatic; Radiologic findings of Ventilator induced lung injury requiring major ventilator adjustments/increased lung protective strategies. | 3 | [8 empty] | 2 | 1 | 1 | 1
VH28 | Contamination | Introduction of debris, dust, or contaminants to patient, user, or environment. | Dual or carrying bag dirty | Infection |  | 1930 | Infection that resolves without intervention (skin infection-see cellulitis) | 1 | none | [7 empty] | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[6 empty] | 1930 | An infection that requires medical intervention Examples: Antibiotics, antifungals, or antivirals, or medication by mouth or inhalation, Incision/drainage see abscess | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 1930 | Symptomatic; IV antibiotics required with increased ventilatory support. An infection that requires more advanced medical intervention | 3 | [8 empty] | 1 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 1 | 1 | 1
VH29 | Damage to internal system components | Internal device system component damage due to oxidation, water ingress, servicing, installation. | Use Error - water ingress | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | N/A | [5 empty] | Caution: Do not reverse the battery polarity when connecting the battery kit to the battery. |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  |  | Use Error - PowerPack is dropped |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | N/A |  |  |  |  | Design - Integrated Quick Blow Fuse |  |  | 1 | 1 | 1 | 1
|  |  | Use Error - polarity of battery is reversed |  | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | N/A |  |  |  |  | Design - PowerPack assembly is designed to have a plastic assembly which would protect the battery packs and the electronics |  |  | 1 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 1 | 1 | 1
VH30 | Device/Device Accessories Become Hot | Any part of device or accessories, where the user handles the device, is hot due to environmental exposure or excessive heating of internal components (e.g. batteries, electrical components). |  | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[6 empty] | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | [8 empty] | 1 | 1 | 1 | 1
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 1 | 1 | 1
VH33 | Electrical shock | The passage of electric current through the body (e.g. exposed electrical components, transmission of electrical current through fluids including bodily). |  | Electrical shock | N | 2554 | Asymptomatic, intervention not indicated. Slight tingling sensation; if burn see Harm of Burn | 1 | N/A |  |  |  |  | Design - Integrated Quick Blow Fuse |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[6 empty] | 2554 | Non-urgent medical intervention indicated. Slight shock felt; not painful but disturbing. Average individual can let go. However, strong involuntary reactions to shocks in this range may lead to injuries .5mA ; if burn see Harm of Burn | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2554 | Medical intervention indicated (planned cardioversion)Painful shock, begin to lose muscular control. Commonly referred to as the freezing current or "let-go" range. 6mA to 16mA.; if burn see Harm of Burn | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2554 | Life-threatening consequences; hemodynamic compromise; urgent intervention indicated; Extreme pain, respiratory arrest, severe muscular contractions. Individual cannot let go. Death is possible. 17mA-99mA; if burn see Harm of Burn | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2554 | Death +G1422 | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Burn, Thermal |  | 2530 | Thermal injury not resulting in need for treatment. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | [8 empty] | 1 | 1 | 1 | 1
VH34 | Environmental specifications exceeded | Ventilator used or stored outside of the environmental specifications as described in the user manual | User Error - environmental temperature or relative humidity exceeded | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | N/A |  |  |  |  | Design - - Battery storage instructions o Do not store above 40deg C. o Batteries must be stored in dry location. o Batteries must be top charged every 3 months. o Never store batteries in sealed environment. |  |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 1 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 1 | 1 | 1
VH38 | Fire Hazard | Fire associated with the device or accessory |  | Burn, Thermal | N | 2530 | Thermal injury not resulting in need for treatment. | 1 | N/A |  |  |  |  | Design - - Battery storage instructions o Do not store above 40deg C. o Batteries must be stored in dry location. o Batteries must be top charged every 3 months. o Never store batteries in sealed environment. |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | 2530 | Minimal repair of the thermal burn using stitches or local wound care. (i.e. May be due to inadvertent activation of the pencil laying on patient or potential internal) | 2 | N/A | [6 empty] | WARNING: Never expose batteries to direct flame. | 1 | 2 | 1 | 1
[5 empty] | N | 2530 | Moderate to major resection is needed to obtain healthy tissue for anastomosis or reconstruction. | 3 | [5 empty] | Design- the battery has a thermal cut off and temperature sensing. The battery will turn off if it overheats. |  |  | 1 | 1 | 1 | 1
[6 empty] | 2530 | Thermal injury not recognized at the time of surgery. Results in the patient having to undergo urgent surgical intervention due to leakage of a structure/organ. (May need to see peritonitis, aroma, sepsis, etc.) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2530 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH39 | Hazardous materials | Device or device materials when discarded may cause effects to patient, user or environment. |  | Chemical exposure, Hazardous | N | 2570 | Minor Inflammation, no treatment needed | 1 | N/A |  |  |  |  | Design- the battery has a thermal cut off and temperature sensing. The battery will turn off if it overheats. |  |  | 1 | 1 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
|  | Improper device disposal |  |  | N | 2570 | Temporary injury or impairment that requires minor professional medical intervention | 2 | N/A |  |  |  |  | Design - sealed battery case |  |  | 1 | 1 | 1 | 1
[5 empty] | N | 2570 | Results in injury or impairment requiring major professional medical intervention; May require hospitalization for treatment due to chemical exposure; Patient stable. | 3 | N/A | [6 empty] | WARNING: Batteries should be disposed of according to environmental legislation in your country and locality. | 1 | 1 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 1 | 1 | 1
VH50 | Inadequate surge protection | Electrical circuit fails during use due to a power surge. | Damaged power cord used | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [5 empty] | Design - Integrated Quick Blow Fuse |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 1 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 2 | 1 | 1
VH54 | Loose Internal mechanical interface or connection | Any internal parts, connectors, or components that become loose or disconnected | Use Error - PowerPack is dropped | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering input and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. The benefits of home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | The battery carrying bag is designed such that the Power Pack display is visible and all connection points are accessible. |  |  | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH55 | Loss of external mechanical interface or connection | Any external parts, connectors, or components that become disconnected | User Error - cables not connected correctly | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 2 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | User Error - Physical damage from impact or vibration |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | The battery carrying bag is designed such that the Power Pack display is visible and all connection points are accessible. |  |  | 2 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [6 empty] | DC DISCONNECTION ALARM |  | 2 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 2 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 2 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 2 | 2 | 1 | 1
VH69 | Inadequate ventilator maintenance | Machine not properly serviced |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 |  |  |  |  | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 2 | 1 | 1
VH72 | No Hazard | No hazard occurred during reported event or identified after event analysis |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [5 empty] | None |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
VH76 | Loose External mechanical interface or connection | Any external parts, connectors, or components that become that become loose. | User Error - Physical damage from impact or vibration | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | The battery carrying bag is designed such that the Power Pack display is visible and all connection points are accessible. |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [6 empty] | DC DISCONNECTION ALARM |  | 1 | 2 | 1 | 1
|  |  |  | No consequences or impact to patient |  | 2199 | Delay in treatment causes user frustration only (Severity would be no harm) | 1 | [8 empty] | 1 | 1 | 1 | 1
VH78 | Ventilator and/or accessories tips, moves or falls | The ventilator or ventilator assembly’s moves, falls or tips on Patient, user, bystander. Impact on environment also possible. | DFMEA PowerPack carry bag Strap failure | Tissue Damage/Tissue Trauma |  | 2104 | Minor tissue damage not requiring medical intervention (minor bruise); Self-limited; Minor bleeding/injury to tissue/vessels that may resolve without medical intervention | 1 | None | [7 empty] | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | Wheelchair attachment fails |  |  | 2104 | Tissue damage requiring minor suturing. Bleeding/injury to tissue/vessels that may resolve with minor medical intervention Examples: minor suture repair may be required (including closing of extended incision), no blood transfusions. | 2 | [8 empty] | 1 | 2 | 1 | 1
|  |  | User drops power pack |  |  | 2104 | Sudden significant injury to tissue that will require professional medical intervention to prevent permanent damage to the tissue | 3 | [8 empty] | 1 | 2 | 1 | 1
|  |  |  | Anxiety |  | 2328 | Mild symptoms; intervention not indicated; eliminate stressors Asymptomatic asynchrony observed on monitoring. | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic; requiring minor change in ventilator settings (e.g., sensitivity, eSens) | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic with vital sign impacts; requiring major change in ventilator settings | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2328 | Symptomatic requiring major change in therapy (e.g., sedation, paralytics, controlled ventilation) | 4 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH81 | EMC Susceptibility | Interference of the ventilators ability to provide ventilation due to electromagnetic energy. |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [5 empty] | none |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  | Note this is looking at the effect on the ventilator if the PowerPack fails due to EMC |  |  | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [5 empty] | Design-Ventilator will transition to internal battery to power the ventilator if the PowerPack fails |  |  | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [7 empty] | WARNING - The Puritan Bennett™ 5x0 requires special precautions for electromagnetic compatibility and should be installed and started according to the recommendations found in Appendix B, “Specifications.” In particular, the use of nearby mobile and portable communications equipment using radio frequencies, such as mobile telephones or other systems exceeding the levels set in the IEC 60601-1-2 standard, may affect its operation. | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [7 empty] | WARNING - The use of any accessory other than those specified, with the exception of the power supplies or cables sold by Covidien, may lead to an increase in electromagnetic emissions or a decrease in the equipment protection against electromagnetic emissions. If the ventilator is used adjacent to such accessories or stacked with such devices, the ventilator’s performance should be monitored to verify normal operation. | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH90 | Trip Hazard | Device or any of its accessories that occupies space within the user's or caregiver's walkway that may lead to tripping | Cables | Fall |  | 1848 | Minor with no resultant injuries; intervention not indicated | 1 | [5 empty] | none |  |  | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
|  |  | Misuse of machine: The user has an incorrect knowledge about the device abilities and intended use. |  |  | 1848 | Symptomatic; noninvasive intervention indicated | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1848 | Hospitalization indicated for injury related to fall; monitoring | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1848 | Urgent surgical treatment indicated due to fall with head injury resulting in a hematoma (See hematoma) needing immediate evacuation to prevent permanent injury. | 4 | [8 empty] | 1 | 1 | 1 | 2
| Use Error - Improper use by non-trained personnel | Non trained personnel monitoring and/or setting patient parameters on the ventilator |  | Respiratory failure |  | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | [8 empty] | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[6 empty] | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 2484 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Dyspnea |  | 1816 | Mild Symptoms, shortness of breath. No treatment required. Self limited. observe | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1816 | Symptomatic; requiring small changes in ventilatory support/oxygen or suction. | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1816 | Deeply Symptomatic; requiring major changes in ventilatory support/oxygen. | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | 1816 | Trouble breathing may be life threatening. Urgent intervention required with intubation or tracheotomy or changing the tube if patient is already intubated. | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | 1816 | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
|  |  |  | Delay of Treatment (Prolonged procedure) |  | N/A | Minor delay, no patient impact | 1 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | [8 empty] | 1 | 2 | 1 | 1
[6 empty] | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | [8 empty] | 1 | 1 | 1 | 2
[6 empty] | N/A | Death | 5 | [8 empty] | 1 | 1 | 1 | 2
VH104 | Battery level indicator is not available | Displayed Battery level indicator is not readable or displayed. |  | Delay of Treatment (Prolonged procedure) | N | N/A | Minor delay, no patient impact | 1 | [7 empty] | The battery charger's LED light will be YELLOW while the battery is charging. When the battery is fully charged, the LED light will switch to GREEN. This transition may be accompanied by a brief flashing of the GREEN LED. | 1 | 2 | 1 | 1 | P2 scores are based on clinical expertise, engineering knowledge and complaint data. | 10038637 PB560-520 Trace Matrix | 10038637 PB560-520 Trace Matrix | Yes | This risk is acceptable and has been reduced as far as possible. Should the power pack fail, there will be transition to the internal battery .The benefits of having a long term battery back up for home care ventilation outweighs this potential risk.
[5 empty] | N | N/A | Delay with minor patient impact (e.g., extended manual ventilation, stable patient). | 2 | PPKPRD29 | [5 empty] | Battery charge indicator LED | There is one LED on the external battery charger. *Orange Start of charge *Yellow Medium level of charge *Green Charge is complete | 1 | 2 | 1 | 1
[5 empty] | N | N/A | Delay with major patient impact (e.g., extended manual ventilation, unstable severe pulmonary status) | 3 | PPKPRD30 | [5 empty] | Battery capacity remaining LED |  | 1 | 1 | 1 | 1
[5 empty] | N | N/A | Prolonged delay with major patient impact,. replacement ventilator unavailable (e.g., extended manual ventilation, unstable severe pulmonary status,) | 4 | PPKPRD31 | [5 empty] | Battery capacity low LED | When the battery is running low, the farthest-left LED on the battery charge indicator turns yellow/orange. | 1 | 1 | 1 | 2
[5 empty] | N | N/A | Death | 5 | PPKPRD97 | [5 empty] | AC mains connected LED |  | 1 | 1 | 1 | 2
|  |  |  | Respiratory failure | N | 2484 | Symptomatic early clinical intervention; Requires increased ventilatory support of one or more parameters | 2 | PPKPRD21 |  |  |  |  | The battery carrying bag shall: - The battery display must be visible through the carrying bag window |  |  | 1 | 1 | 1 | 1
[5 empty] | N | 2484 | Symptomatic; Requires multiple ventilator adjustments, administration of medication, and/or elective change from non-invasive to invasive ventilation. | 3 | PPKPRD25 |  |  |  |  | The battery carrying bag is designed such that the Power Pack display is visible and all connection points are accessible. |  |  | 1 | 1 | 1 | 1
[6 empty] | 2484 | Symptomatic; Requires move from non-invasive to invasive ventilation with emergent intubation. | 4 | [8 empty] | 1 | 1 | 1 | 2

## Drop Downs (hidden)
Risk Control Option | E P | Acceptability | Severity | Occurrence
Inherently Safe By Design | Yes | Yes | 1 | 1
Protective Measures | No | No | 2 | 2
Information For Safety |  |  | 3 | 3
No Further Action |  |  | 4 | 4
|  |  | 5 | 5
|  | Hazard Status |  | Attendance
|  | Preliminary | Risk Index | P
11 |  | Confirmed | 3 | A
12 |  | Analyzed | 2 | C
13 |  | Mitigated | 1 | E
14 |  | Verified
15 |  | Closed
21 |  | Closed as not credible
22
23
24
25
31
32
33
34
35
41
42
43
44
45
51
52
53
54
55