## PFMEA Summary
Change Summary:
Revision | Date | Author | Description
[redacted]
Summary:
This Risk process Management covers the PB560 Ventilator. It was completed with input from the following team: .
Access Residual Risk - RPN # after Risk Mitigations have been implemented (Refer to Process Failure Analysis Procedure DC00070840:

## Process Flow Map
Process Flow Map for PB560
Process: | PB560 Assembly |  | Prepared by | [redacted] |  | Activity Classes | Operation | Delay
Responsibility | [redacted] |  | Date: | Refer to Agile |  |  | Inspection | Transport
|  |  | Revision | U |  |  | Inspection/Operation | Storage
Decision
ORIGIN | INPUTS |  |  | PROCESS STEPS |  |  | OUTPUT |  |  | Destination
| KPIV's | Msm Sys | Type | Activity Class | Process Step |  | KPOV's | MSN sys | Specs
Receiving
Start | Receiving, | Part Numbers, Revisions | Procedure | Operation, Storage |  |  | New Material for disposition | Galway Requirements Planning System (BPCS) | as per Galway Requirements Planning System (BPCS) | Warehouse Storage
Receiving | Incoming Inspection, Released Mfg drawings, IPC-A-610, II019105 | Visual, verniers, Optical. | Equipment | Inspection |  | Incoming Inspection | Acceptable material | Released Mfg drawings, | as per drawing | Line storage
Incoming Inspection, Stores | Unpack material & store in tote bins etc | N/A | Procedure | Storage |  | Tube Cutting | Segregated material | BOMs for product |  | Tube Cutting
Storage location | Fixtures | Calibrated Ruler | Equipment | Operation | [5 empty] | Lower Enclusore Assembly
Storage location | Assembly Procedure 10021625, PHR 10021626 | G-QAP004 | Procedure | Operation | [5 empty] | Lower Enclusore Assembly
Storage location | Materials | 10021625 | Material | Operation |  |  | Material cut to length | Calibrated ruler | 10095161, 10095162 & 10095163 | Lower Enclusore Assembly
Storage location | Trained operator |  | Procedure | Operation | [5 empty] | Lower Enclusore Assembly
Storage location | Unpack material & store in tote bins etc | N/A | Procedure | Storage |  |  | Segregated material | BOMs for product
Storage location | Fixtures | 3834400, 3835100, 3834500, 3835000 | Equipment | Operation |  | Sub Assembly, Main Assembly. | Assembled unit | 10095161, 10095162 & 10095163 |  | Final Assembly Enclosure
Storage location | Assembly Procedure 10095161, 10095162 & 10095163 PHR 10038461 | G-QAP004 | Procedure | Operation |  |  | Bottom Enclosure sub-assembly |  |  | Top Enclosure Assembly
Storage location | Materials | 10095161, 10095162 & 10095163, 4096600 | Materials | Operation | [5 empty] | Top Enclosure Assembly
Storage location | Air De-ioniser and air gun | Charge Plate Analyser | Equipment | Operation |  |  |  | G-QCAL041 and G-QCAL234 |  | Top Enclosure Assembly
Storage location | Torque tooling | Static Torque setter, URYU | Tooling | Operation |  |  | Torque set to 4.4in-lb, 8.9in-lb | G-QCAL030 | 8.9in-lb +/- 0.5, 4.4in-lb +/- 0.5, | Top Enclosure Assembly
| Trained operator |  | Procedure | Operation | [5 empty] | Top Enclosure Assembly
Storage location | Assembly Procedures 10095161, 10095162, 10095163 |  | Procedure | Operation |  |  | Assembled unit |  |  | Safety Analyzer Test
Storage location, Bottom Enclosure Assy | Materials | 10095161, 10095162 & 10095163 4096600 | Materials | Operation |  |  | Assembled unit |  |  | Safety Analyzer Test
Storage location | Torque tooling | Static Torque setter, URYU | Tooling | Operation |  |  | Torque set to 4.4in-lb, 8.9in-lb | G-QCAL030 | 8.9in-lb +/- 0.5, 4.4in-lb +/- 0.5, | Safety Analyzer Test
| Trained operator | G-QAP017 | Procedure | Operation | [5 empty] | Safety Analyzer Test
Storage location | Air De-ioniser and air gun | Charge Plate Analyser | Tooling | Operation |  |  |  | G-QCAL041 and G-QCAL234 |  | Safety Analyzer Test
Safety Test - Safety Analyzer
Storage location | Metron QA-90 Safety Test Analyser |  | Equipment | Operation |  |  | Calibrated Tester | 10036007
Storage location | Cables |  | Equipment | Operation |  |  |  | 10036007 |  | Download Process
Storage location | UUT |  | Equipment | Operation |  |  | Unit with application s/w | 10036007 |  | Download Process
Storage location | Safety Test Process |  | Equipment | Operation |  |  |  | 10036007 |  | Download Process
| Trained Operator | 10036007 | Procedure |  |  |  | Unit with application s/w | 10036007 |  | Download Process
Storage location | Cables |  | Equipment | Operation |  | Software Download Process |  |  |  | System Test
Storage location | PC |  | Equipment | Operation |  |  | Unit with application s/w | 10036007 |  | System Test
Storage location | UUT |  | Equipment | Operation |  |  | Unit with application s/w | 10036007 |  | System Test
Storage location | S/W |  | Equipment | Operation | [5 empty] | System Test
Storage location | Download process |  | Equipment | Operation |  |  |  | 10036007 |  | System Test
| Trained Operator | 10036007 | Procedure |  |  |  | Unit with application s/w | 10036007 |  | System Test
System Calibration & Test
Storage location | System Test procedure | 10036007 | Procedure | Inspection /Test |  |  | Unit with s/w downloaded and calibration tested | 10036007 |  | Test & Battery Charge
Storage location | Trained Operator | 10036007 | Procedure | Inspection /Test |  |  | Unit with s/w downloaded and calibration tested | 10036007 |  | Test & Battery Charge
Storage location | PC & S/W |  | Equipment | Inspection /Test |  |  | Unit with s/w downloaded and calibration tested | 10036007 |  | Test & Battery Charge
Storage location | UUT |  | Equipment | Inspection /Test |  |  | Unit with s/w downloaded and calibration tested | 10036007 |  | Test & Battery Charge
Storage location | Cables |  | Equipment | Inspection /Test
Storage location | Assembly Procedure | 10036007 | Proceduure | Operation / Test |  | Battery Discharge / Charge | Unit ready for Final System Test | 10036007 |  | System Final Test
Storage location | Materials | 4096600 | Materials | Operation / Test |  |  | Unit ready for Final System Test | 10036007 |  | System Final Test
| Trained operator | 10036007 | Procedure | Operation |  |  | Unit ready for Final System Test | 10036007 |  | System Final Test
Storage location | Test Equipment |  | Equipment | Operation |  |  | Unit ready for Final System Test | 10036007 |  | System Final Test
Storage location | System Test procedure | 10039030 | Procedure | Inspection /Test |  | System Final Test | Final Tested Unit | 10039030 |  | Pack Unit
Storage location | Trained Operator | 10039030 | Procedure | Inspection /Test |  |  | Final Tested Unit | 10039030 |  | Pack Unit
Storage location | PC & S/W | 10039030 | Procedure | Inspection /Test |  |  | Final Tested Unit | 10039030 |  | Pack Unit
Storage location | UUT |  | Equipment | Inspection /Test |  |  | Final Tested Unit | 10039030 |  | Pack Unit
Storage location | Cables |  | Equipment | Inspection /Test |  |  | Final Tested Unit | 10039030 | PB560 10039196 | Pack Unit
Storage location | Assembly Instruction | 10039897 | Procedure | Operation /Inspection |  | Pack unit (All product) | Packed Unit | 10039897 |  | Product Release Authorisation
Storage location | Materials |  | Materials | Operation |  |  |  | 10039897 |  | Product Release Authorisation
On line Equipment | UDI Barcode Grading Verifier / equipment | QS00002016 | procedure | Operation / Inspection |  |  | Device & Carton Labels Verified as UDI Compliant | G-QAP019 (Calibration procedure) | Tasks performed as per applicable procedures | PB560 Packing & Labelling Procedure
Workstation | Verification Scanner | PM: Testers and Test Stands, G-AMFG-1180-00,10149038 verification scanner operation procedure | Procedure | Operation/Inspection |  |  | Packed Unit | 10149038 |  | Product Release
| Trained Operator | 10039897 | Procedure | Operation |  |  |  | 10039897 |  | Product Release Authorisation
Storage location | UUT |  | Equipment | Operation | [5 empty] | Product Release Authorisation
Product Audit Release Authorisation (All product)
Storage location | Assembly instruction | 10021754 | Procedure | Operation |  |  | Product Released for Shipment | Signature on DHR 10038461 |  | Shipping
Storage location | Trained Operator | 10021754 | Procedure | Operation |  |  | Product Released for Shipment | Signature on DHR 10048361 |  | Shipping
Ship unit
Customer
| Humidity | Humidity Sensor | Procedure | Inspection |  | Environmental | Product manufactured in acceptable humidity range | Humidity monitored per G-AMFG-1110-00 | 22%-78% RH
| Temperature | Hanwell Temperature | Procedure | Inspection |  |  | Product manufactured in acceptable temperature range | Temperature monitored per G-AMFG-1110-00 | 16.5 to 29.5deg C
10037310U00

## FMEA
Process or Product Name: |  |  |  |  | PB560 Assembly, Test & Pack | [5 empty] | Updated by: [redacted] |  |  | Page __1__ of __16__
Responsible: [redacted] | [10 empty] | FMEA Date (Refer to AGILE) | U
CTQ | CTS | Process Function | Process Step | Key Process Input | Potential Failure Mode | Potential Failure Effects | Hazard Tag | Severity | Potential Causes | OCC | Current Controls | DET | RPN | Actions Recommended | Resp. | Actions Taken | SEV | OCC | DET | RPN
Critical to Quality | Critical to Safety: Risk documents listed in this column must be reviewed prior to making any changes to procedures in Key Process Input or Current Controls | Operation Name for Manufacturing task | Description of step within operation | What is the Key Process Input? | In what ways does the Key Input go wrong? | What is the impact on the Key Output Variables (Customer Requirements) or internal requirements? | Unique Identifier for the hazard | How severity effects the process | What causes the Key Input to go wrong? | How often does cause or FM occur? | What are the existing controls and procedures (inspection and test) that prevent either the cause or the Failure Mode? Should include an SOP number. | How well can you detect cause or FM? |  | What are the actions for reducing the occurrence of the Cause, or improving detection? Should have actions only on high RPN's or easy fixes. | Whose Responsible for the recommended action? | What are the completed actions taken with the recalculated RPN? Be sure to include completion month/year
|  | Tube Cutting /Assembly (PB560 Assy Procedure 10095162 / 10021625) | Tube Cutting /Assembly (PB560 Assy Procedure 10095162 / 10021625) | Tube cutting fixture set up. | Incorrect setting | Tubes cut to the incorrect length, can not assemble | N/A | 4 | Procedure not followed, operator error | 10 | Measure at start and end per procedure 10095162; Ruler is calibrated and metric only. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tube cutting fixture set up. | Incorrect setting | Tubes cut to the incorrect length causing kinks. | N/A | 4 | Procedure not followed, operator error | 10 | Measure at start and end per procedure 10095162;; Ruler is calibrated and metric only. Tubes selected prior to assembly onto valve. Product 100% tested. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tube cutting | Incorrect setting | Tube cut shorter than specification but can be assembled | N/A | 7 | Procedure not followed, operator error | 10 | Measure at start and end per procedure10095162;; Ruler is calibrated and metric only. Tubes selected prior to assembly onto valve. Product 100% tested. | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tube cutting | Flash after cut | Potentially impact on sensor | N/A | 7 | poor quality cut | 1 | Precision fixture. Inspection of first and last part during cutting for flash. 540 PM procedure verifys clean cut. | 3 | 21 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tubing specification | Incorrect Tubing type or size | Tubes do not assemble to fittings or is incorrect type | N/A | 7 | Miss-labelling of tubing | 10 | Supplier Qualification: Receiving against manufacturer's p/n. Incoming Inspection Process. | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tubing specification | Incorrect Tubing type or size | Tubing kinks. | N/A | 7 | Miss-labelling of tubing | 10 | Supplier Qualification: Receiving against manufacturer's p/n. Incoming Inspection Process. | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Dirt entering tubing | Impact performance of assembly | N/A | 7 | Dirt in production area | 1 | Production area cleaned daily | 1 | 7 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Silicone Bumpers Missing | Potential reliability issue in the field | N/A | 4 | Process operative error | 2 | Visual, Cosmetic Check, Final QA. | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  | Inspiratory Conical Block (PB560 Assy Procedure 10095162) | Inspiratory Conical Block (PB560 Assy Procedure 10095162) | Trained Operator | Label placed with air bubble | Cosmetic Impact | N/A | 2 | Process operative error | 10 | Visual, Cosmetic Check, Final QA. | 3 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Label incorrectly placed | Customer is miss directed | N/A | 8 | Process operative error | 10 | Training; all procedures will be peer reviewed prior to approval; label shape is unique to application. Procedure contains photo checks | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Inspiratory block is not cleaned | Label falls off | N/A | 8 | Process operative error | 10 | Process instruction; training. Final inspection contains a label check. | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Inspiratory block | Flash on tube fittings | Leaks - obstructions | N/A | 7 | Vendor quality issue | 10 | Qualified Vendor / Galway Incoming Inspection Procedure.100% inspection in place at incoming | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tubes are incorrectly placed | Tubing does not assemble | N/A | 4 | Process operative error | 10 | If tubings misoriented, assy will not assemble into unit (diff. lengths of tubing) | 2 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tubes are incorrectly placed | Kinks in the tubing. | N/A | 5 | Process operative error | 10 | 10095162;, training. Photos in procedure for direction. | 1 | 50 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | FiO2 connector is incorrectly aligned with inspiratory block | Connector cannot be tightened correctly. | N/A | 2 | Process operative error | 10 | Parts are keyed, 10095162;, training | 3 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | FiO2 connector is not tightened fully as connector is close to side of inspiratory block. Fixture torques against inspiratory block. | Connector is loose | N/A | 4 | Fixture 10038462 contacts inspirotory block wall while tightening | 10 | Training. Additional warnings in procedure. Fitting will be loose during test. | 2 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Dirt entering conical block | Impact performance of assembly | N/A | 7 | Dirt in production area | 1 | Production area cleaned daily | 1 | 7 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Cable harness not routed correctly / too much pressure. | Cable makes loose electrical connection over time leading to failure | N/A | 7 | Process operative error | 2 | Process instruction training. | 3 | 42 | N/A | N/A | N/A | - | - | - | -
|  | Expiratory Conical Block (PB560 Assy Procedure 10095162) | Expiratory Conical Block (PB560 Assy Procedure 10095162;) | Trained Operator | Label placed with air bubble | Cosmetic Impact | N/A | 2 | Process operative error | 10 | Visual, Cosmetic Check, Final QA. | 3 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Label incorrectly placed | Customer is miss directed | N/A | 8 | Process operative error | 10 | Training; all procedures will be peer reviewed prior to approval; label shape is unique to application. Procedure contains photo checks | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Expiratory block is not cleaned | Label falls off | N/A | 8 | Process operative error | 10 | Process instruction; training. Final inspection contains a label check. | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tubes are incorrectly placed | Tubing does not assemble | N/A | 4 | Process operative error | 10 | If tubings misoriented, assy will not assemble into unit (diff. lengths of tubing) | 2 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tubes are incorrectly placed | Kinks in the tubing. | N/A | 5 | Process operative error | 10 | 10095162;, training. Photos in procedure for direction. | 1 | 50 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tubes are loosely connected | Tube becomes disconnected over time | N/A | 8 | Process operative error | 10 | process instruction 10095162;training; test instructions 10036007; 10039030. Check before closing unit. | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Driver set below specification | Fittings become loose | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095162; | 5 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Driver set above specification | Fittings are damaged | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095162; | 2 | 16 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Dirt entering exhalation block | Impact performance of assembly | N/A | 7 | Dirt in production area | 1 | Production area cleaned daily | 1 | 7 | N/A | N/A | N/A | - | - | - | -
|  | Battery Cover (PB560 Assy Procedure 10095163) | Battery Cover (PB560 Assy Procedure 10095163) | Process task | Label placed with air bubble | Cosmetic Impact | N/A | 2 | Process operative error | 10 | Training,10095163; Cosmetic check in procedure | 3 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Battery Cover is not cleaned inside | Battery label falls off | N/A | 4 | Process operative error | 10 | Process instruction; training. Final inspection contains a label check. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Battery Cover is not cleaned inside | Adhesive stops fall off, battery is not secured. | N/A | 7 | Process operative error | 10 | Training, 10095163; Only potential opportunity for stops to fall off is during assembly; Stops checked immeadiately prior to assembly | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Adhesive stops not fitted | Battery in not secured, system connection to battery could be damaged | N/A | 7 | Process operative error | 10 | Training, 10095163 Only potential opportunity for stops to fall off is during assembly; Stops checked immeadiately prior to assembly | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Battery Cover is not cleaned outside | Label falls off, no direction for user. | N/A | 8 | Process operative error | 10 | Shipping audit. Customer has the operation manual for reference. There is a label check at pack. | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  | Blower Housing (PB560 Assy Procedure 10095162) | Blower Housing (PB560 Assy Procedure 10095162) | Process task | Foam strips too long | Interference to overall assembly | N/A | 5 | Process operative error | 10 | Specific length per procedure. Cut with calibrated fixture. Length is verified by how it fits in assembly. | 1 | 50 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Foam strips too short | Blower assembly could vibrate | N/A | 3 | Process operative error | 10 | Specific length per procedure. Cut with calibrated fixture. Length is verified by how it fits in assembly. | 1 | 30 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Turbine not cleaned | Foam on top of blower might not stick. | N/A | 4 | Process operative error | 10 | Training, 10095162 (top of blower is cleaned); Foam will be held in place by top housing. Occurrence of falling off are remote. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Turbine not cleaned | Collar does not stick and Power Supply / Blower Harness ferrite is unsupported | N/A | 4 | Process operative error | 10 | Training, 10095162 (top of blower is cleaned). Check in procedure. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Spacers too tight | Turbine housing damage | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095162 | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Spacers too loose | Spacers work loose | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095162 | 5 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Loctite not applied | Screws work loose | N/A | 8 | Process operative error | 10 | Training, 10095162. Procedure contains check for loctite | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Blower card mis-orientated | Can not assemble shield | N/A | 4 | Process operative error | 10 | Training, 10095162 | 2 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Cable is not wrapped correctly | Cable is pinched between shield and turbine. | N/A | 7 | Process operative error | 10 | Training, 10095162(check in place) | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Mis-handling of blower card - ESD damage | Unit fails over time; impacts production yield | N/A | 8 | Process operative error | 10 | Training | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Mis-handling of blower card | Physical damage | N/A | 8 | Process operative error | 10 | Training | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Connect the power supply/blower harnesses (3822500) incorrectly | Blower does not function | N/A | 8 | Process operative error | 10 | 10095162 procedure, training, 10036007 / 10039030 test procedures, Keying of connection | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | power supply/blower Harnesses not connected | Blower does not function | N/A | 8 | Process operative error | 10 | 10095162 procedure, training, 10036007 / 10039030 test procedures, Keying of connection | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tubing not connected | Blower does not function | N/A | 8 | Process operative error | 10 | 10095162 procedure, training, 10036007 / 10039030 test procedures, Keying of connection | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tubing specification | Incorrect Tubing type or size | Tubes do not assemble to fittings. | N/A | 4 | Miss-labelling of tubing | 3 | Control of suppliers (AML) ref G-QAP007; GMP audit to be completed as part of validation (material is segregated) | 1 | 12 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tubing specification | Incorrect Tubing type or size | Tubing kinks. | N/A | 7 | Miss-labelling of tubing | 3 | Control of suppliers (AML) ref G-QAP007; GMP audit to be completed as part of validation (material is segregated) | 3 | 63 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tubing connected to incorrect fitting | Will fail test | N/A | 8 | Process operative error | 10 | 10036007 / 10039030 test procedures | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Hand Torque Driver | Torque too high | Damage threading on blower housing | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095162; If torque is too high, threading may be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 21 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Hand torque Driver | Torque too low | Collar could shake loose | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095162 | 6 | 42 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Collar not fitted | Power Supply Blower Harness ferrite is unsupported | N/A | 4 | Process operative error | 10 | Note in procedure to check for collars | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | CPU / Power supply harness (3822300) not connected | Blower does not function | N/A | 8 | Process operative error | 10 | 10095162 procedure, training, 10036007 / 10039030 test procedures. | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | CPU / Power supply harness (3822300) not assembled correctly | Blower does not function | N/A | 8 | Process operative error | 10 | 10095162 procedure, training, 10036007 / 10039030 test procedures, Keying of connection | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | CPU / blower driver harness (3822400) not connected | Blower does not function | N/A | 8 | Process operative error | 10 | 10095162 procedure, training, 10036007 / 10039030 test procedures. | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | CPU / blower driver harness (3822400) not assembled correctly | Blower does not function | N/A | 8 | Process operative error | 10 | 10095162 procedure, training, 10036007 / 10039030 test procedures, Keying of connection | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Dirt entering blower | Impact performance of assembly | N/A | 7 | Dirt in production area | 1 | Production area cleaned daily | 1 | 7 | N/A | N/A | N/A | - | - | - | -
|  | Upper Housing (PB560 Assy Procedure 10095163) | Upper Housing (PB560 Assy Procedure 10095163) | Process task | Handle mis-orientated | Parts will not assemble | N/A | 4 | Process operative error | 10 | Only one orientation, 10095163 procedure, training | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Screws too tight | Handle is damaged | N/A | 4 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095163 | 6 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Screws too loose | Handle becomes loose | N/A | 4 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095163 | 6 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Housing is not cleaned | Membrane does not stick down properly | N/A | 7 | Supplier neglects to clean housing / Process operator neglects to clean housing | 3 | Supplier cleans Housing prior to painting process: Operator cleaning task is secondary. Training, procedure 10095163 | 3 | 63 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Housing is not cleaned | Alarm Label does not stick down properly | N/A | 4 | Supplier neglects to clean housing / Process operator neglects to clean housing | 3 | Supplier cleans Housing prior to painting process: Operator cleaning task is secondary. Training, procedure 10095163 | 3 | 36 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Membrane not located correctly | Damage due to relocation | N/A | 7 | Process operative error | 10 | Design of membrane and corresponding recess in top housing: only one possible install. | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Membrane not located correctly | Cosmetic Impact | N/A | 4 | Process operative error | 10 | Training. References in procedure | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Flex tail not located through slot | Part will not assemble | N/A | 4 | Process operative error | 10 | 10095163; training | 2 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Even pressure is not applied | Membrane will become detached | N/A | 7 | Process operative error | 10 | Training. References in procedure | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Switch panel damaged during manufacture or assembly. | Push button or Alarm LED not working | N/A | 8 | Damage due to mis handling during manufacture, shipping, assembly. | 5 | Physical condition checked as part of IIP, system test excercises all keypad buttons and LEDs. Keypads lot controlled. Warnings in mainline procedure regarding handling of keypads. | 3 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review failure rate for keypads to revise occurance. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data for PB540. Based on this data, reduce occurrence | 8 | 3 | 3 | 72
|  | Switch Cover (PB560 Assy Procedure 10095161) | Switch Cover (PB560 Assy Procedure 10095161) | Process task | Spring mis-assembled | Spring does not work | N/A | 7 | Process operative error | 10 | Check in procedure 10095161 to ensure correct operation | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Axle not cantered | Parts will not assemble | N/A | 4 | Process operative error | 10 | Check in procedure 10051246 to ensure correct operation | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Place spring through the wrong opening | Parts will not assemble | N/A | 4 | Process operative error | 10 | Check in procedure 10051246 to ensure correct operation | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Retaining ring not secured correctly to lower housing | Parts will become dislodged in the field | N/A | 8 | Part not assembled correctly as per procedure | 10 | Check in procedure 10095161 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  | Lower Housing (PB560 Assy Procedure 10095161) | Lower Housing (PB560 Assy Procedure 10095161) | Process task | ESD label not fitted | Service warning not present | N/A | 2 | Process operative error | 10 | All Product should be serviced in field only be trained service personnel using Service manual. | 3 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Air inlet label not fitted | End user will may need to refer to manual | N/A | 4 | Process operative error | 10 | Revision of label needs to be recorded on DHR as per 10095161. Product Release procedure verifies DHR completion. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | O2/RA/SP label not fitted | End user will may need to refer to manual | N/A | 4 | Process operative error | 10 | Revision of label needs to be recorded on DHR as per 10095161. Product Release procedure verifies DHR completion. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Identification label is not fitted | Traceability is lost on the device. | N/A | 8 | Process operative error | 10 | Revision of label needs to be recorded on DHR as per 10095161 S/N is read during several steps of process | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Warning label is not fitted | Operator will be mis-directed | N/A | 6 | Process operative error | 10 | Revision of label needs to be recorded on DHR as per 10095161. Product Release procedure verifies DHR completion. | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Adhesive feet not fitted. | Unit can slide on resting surface / potential excess vibration on some surfaces | N/A | 1 | Process operative error | 10 | Procedure 10095161 training | 1 | 10 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Polymide Tube Prep Tool | Tube 'Flaring' Tool not employed | Ends of o2 Connector Tubing are not widened, making insertion more difficult. | N/A | 2 | Process operative error | 10 | Detail added to Procedure 10095161 to use fixture prior to insertion, also statement "ensure tubing is pushed all the way in". Operators are Trained in procedure and in use of fixture. | 3 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Over tighten O2 connection | Damage to connection | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161; If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 21 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Under tighten O2 connection | O2 connection will come loose | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161, 10095162, 10095163; connection is used during test process; unit is serviced periodically in the field; | 4 | 28 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Over tighten O2 connection fastener | Connector damage | N/A | 3 | Process operative error | 10 | Procedure 10095161, 10095162, 10095163; training | 1 | 30 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Under tighten O2 connection fastener | Leaks | N/A | 7 | Process operative error | 10 | Procedure 10095161, 10095162, 10095163; training | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Over tighten alarm call connection | Damage to connection | N/A | 3 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161; If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 6 | 18 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Under tighten alarm call connection | Alarm call connection will come loose | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161; connection is used during test process; unit is serviced periodically in the field; | 6 | 42 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Knurled knob not fitted | Expiratory block can be installed but cannot be secured | N/A | 5 | Process operative error | 10 | Procedure 10095161; training. Check at final QA | 1 | 50 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Foam is cut too short | Turbine will not sit level if the foam moves during assembly. | N/A | 3 | Process operative error | 10 | Training, 10095161 | 1 | 30 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Foam is cut too long | Turbine will not fit correctly and will not assemble | N/A | 3 | Process operative error | 10 | Training, 10095161 | 1 | 30 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Foam not placed | Turbine will not be secured in place | N/A | 3 | Process operative error | 10 | Training; check in place in 10095161 to ensure that foam is in place. Foam is still visible after blower is integrated into assembly | 1 | 30 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too high | Damage to O2 valve (due to 4 screws being overtightened) | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161; If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 21 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too low | O2 valve could shake loose (due to 4 screws being undertightened) | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161 | 6 | 42 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Bracket fitted wrong way around | Cause assembly issues later | N/A | 4 | Process operative error | 10 | Procedure w/ 3 pictorials | 2 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tubing Specification | Incorrect type of tubing used | Assembly issues | N/A | 4 | Miss-labelling of tubing tote bins | 3 | Receiving; p/n identification; | 1 | 12 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tubing Specification | Incorrect length of tubing | Assembly issues | N/A | 4 | Miss-labelling of tubing tote bins | 3 | GMP, 10095161, 10021625 | 1 | 12 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Tubing Specification | Incorrect length of tubing | Kinking | N/A | 7 | Miss-labelling of tubing tote bins | 3 | Training, 10095161 GMP | 1 | 21 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tube put on incorrect connection | O2 mixing will not function | N/A | 8 | Process operative error | 10 | Training, 10095161 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Ty Wrap not fitted | O2 leaks | N/A | 7 | Process operative error | 10 | Training, 10095161. Check added to procedure | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Tube placed on incorrect connection | O2 mixing will not function | N/A | 8 | Process operative error | 10 | Training, 10095161 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Threaded cap not fitted | O2 leaks | N/A | 7 | Process operative error | 10 | Training, 10095161 Check added to procedure | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Collar / Ferrite | Wiring pinched/crushed | N/A | 7 | Process operative error | 10 | Training, 10095161. Check added to procedure | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  | Lower Housing II (PB560 Assy Procedure 10095161) | Lower Housing II (PB560 Assy Procedure 10095161) | Process task | Storage of fans | Fan balancing compromised | N/A | 8 | Storage inappropriate | 2 | Stored unsegregated, in tote | 1 | 16 | Totes with compartments to be sourced | N/A | N/A | - | - | - | -
|  |  |  | Process task | Fan positioned upside down | Fan will not function as efficiently as it should. | N/A | 5 | Process operative error | 10 | Training, 10095161, 10038461. Check added to procedure | 2 | 100 | ALARP - Review PFMEA as part of G-QAP056. Review failure rate for internal temperature and customer complaint rate for evidence of occurance. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 5 | 4 | 2 | 40
|  |  |  | Process task | Fan positioned with cable exiting incorrectly | Will not assemble | N/A | 8 | Process operative error | 10 | Training, 10095161 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | kinked cable | noise | N/A | 4 | incorrect installation | 2 | Operator training storage | 2 | 16 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too high | Lower enclosure threading will not torque | N/A | 4 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161; If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 12 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too low | Fan could shake loose | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161. Check in procedure | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Enclosure cracks where screw is inserted. | Fan could shake loose | N/A | 8 | Screw Hole in enclosure is incorrectly sized. | 2 | Check for crack added to mainline procedure 10095161 | 3 | 48 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Card mis-orientated | Will not assemble | N/A | 4 | Process operative error | 10 | Training, 10095161 | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too high | Damage to card stop | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161; If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too low | Card stop could shake loose | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161. Check in procedure | 6 | 48 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not insert the power card rear first | Will not assemble | N/A | 4 | Process operative error | 10 | Training, 10095161 | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not attach fan | Unit will run hot | N/A | 8 | Process operative error | 10 | Product SW will detect if fan is not connected "COOLING FAN FAILURE"; Training,10095161, 10038461. DHR has specific check. | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Handling of PCBA | Warpage of Power supply PCBA board | component failure | N/A | 8 | poor handling/tool slippage dropping of the board | 1 | Operator training storage | 2 | 16 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
|  |  |  | Torque Driver | Torque too high | Damage to power card (PCBA) | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161; If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too low | Power card (PCBA) could shake loose | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161; connection is used during test process; unit is serviced periodically in the field; | 4 | 28 | N/A | N/A | N/A | - | - | - | -
|  | Final Assembly (PB560 Assy Procedure 10095163) | Final Assembly (PB560 Assy Procedure 10095163) | Process task | Do not connect silicone tubing | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030. Specific checks in procedure | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect Power Supply Blower Harness fully | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect tubing from O2 valve fully | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect cable tie | O2 leaks | N/A | 5 | Process operative error | 10 | Training, 10095163 Check on DHR | 1 | 50 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect cable tie | O2 tube becomes disconnected. Unit will fail. | N/A | 4 | Process operative error | 10 | Training, 10095163; 10036007, 10039030 | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Catch cables with blower assembly | Damage wiring | N/A | 8 | Process operative error | 10 | CAUTION Note contained in PI 10095163. | 2 | 160 | ALARP - Review PFMEA and test procedure to determine if RPN score can be reduced further. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | N/A | - | - | - | -
|  |  |  | Process task | Do not connect cables to buzzer card fully | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect the tubing correctly | Unit will fail test or come loose in field over time | N/A | 8 | Process operative error | 10 | Training, 10051250; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not fit ferrite to FiO2 cable | Unit could be succeptable to emissions. | N/A | 8 | Process operative error | 10 | Training, 10095163 | 5 | 400 | ALARP - Change sequence of assembly to include ferrite during sub assembly. Add check to DHR. | Sean Knight | Updated mainline procedure. Ref ECO-R180705 | 8 | 5 | 1 | 40
|  |  |  | Process task | Pinch FiO2 cable in ferrite | Unit will not have FiO2 functionality | N/A | 5 | Process operative error | 10 | Training,10095163; 10036007, 10039030 | 1 | 50 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | FiO2 wire is not wound around ferrite | Unit could be succeptable to emissions. | N/A | 8 | Process operative error | 10 | Training, 10095163 | 5 | 400 | ALARP - Change sequence of assembly to include ferrite during sub assembly. Add instruction regarding routing cables. | Sean Knight | Updated mainline procedure. Ref ECO-R180705 | 8 | 5 | 1 | 40
|  |  |  | Process task | FiO2 wire is left too long between J5 and ferrite | Unit could be succeptable to emissions. | N/A | 8 | Process operative error | 10 | Training, 10095163 | 5 | 400 | ALARP - Change sequence of assembly to include ferrite during sub assembly | Sean Knight | Updated mainline procedure. Ref ECO-R180705 | 8 | 5 | 1 | 40
|  |  |  | Process task | FiO2 wire is left too short between J5 and ferrite | Connection J5 cannot be made | N/A | 4 | Process operative error | 10 | Training, 10095161 | 1 | 40 | N/A | N/A | N/A
|  |  |  | Process task | Tubing gets pinched | Unit will fail test | N/A | 7 | Process operative error | 10 | Training, 10095163 10036007, 10039030 | 1 | 70 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too high | Damage to card. | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161, 10095162, 10095163; If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too low | Card could shake loose | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161, 10095162, 10095163; connection is used during test process; unit is serviced periodically in the field; | 4 | 28 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect O2 valve | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect SPO2 valve | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect alarm repeater harness | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect buzzer connection | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect power cord | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect blower connection. | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10051250; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Put foam in backwards | Harder to see if the filter is dirty | N/A | 3 | Process operative error | 10 | PI specifies to mount black side of foam toward inside of unit. Specific check in procedure | 1 | 30 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not connect keypad | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not lock keypad flex tail connection | Unit will operate intermittently | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030. Specific checks in procedure | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Pinch keypad cable in enclosure | Prevent functionality of keypad / keyboard. Fail test. Keyboard Flex Tail Track(s) can be broken / damaged. | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too high | Damage to enclosure. | N/A | 6 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161,10095162 & 10095163. If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 18 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too low | Enclosure could shake loose | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161,10095162 & 10095163; connection is used during test process; unit is serviced periodically in the field; | 4 | 28 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Loose screw left in vent assembly. | Screw could short an electrical circuit. | N/A | 8 | Operator accidently drops screw from screwdriver. | 10 | Training, 10095163, screws are inserted one at a time. | 1 | 80 | N/A | N/A | N/A
|  |  |  | Process task | Screw becomes loose and free to move within unit. | Screw could short an electrical circuit. | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095163 | 6 | 48 | N/A | N/A | N/A
|  | Mount the Battery (PB560 Assy Procedure 10051250 | Mount the Battery (PB560 Assy Procedure 10051250 | Process task | Insert battery incorrectly | Unit will fail test | N/A | 8 | Process operative error | 10 | Training, 10095163; 10036007, 10039030 | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too high | Damage to battery cover. | N/A | 8 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161,10095162 & 10095163; If torque is too high, plastic threading will be shredded and torque tool will not 'click' (indication of torque level achieved). | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Torque Driver | Torque too low | Cover could shake loose | N/A | 7 | Torque driver is incorrect set. | 1 | Calibration Procedure; if tool dropped, CAL Lab are notified etc as per 10095161,10095162 & 10095163; connection is used during test process; unit is serviced periodically in the field; | 4 | 28 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Do not fit the cable restraint. | Mains cable could become disconnected. | N/A | 4 | Process operative error | 10 | Training, 10095163 | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  | Hipot Test | Hipot Test | PB500 ELECTRICAL SAFETY TEST & BUZZER CARD CHARGE PROCEDURE - 10076002 | Procedures misinterpreted due lack of clarity; incorrectly written | Hipot test is not executed correctly or unsafely | N/A | 9 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 1 | 90 | ALARP - Review PFMEA and test procedure as part of G-QAP056. Review procedure vs practice. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 9 | 2 | 1 | 18
|  |  |  | Operator: | Operator: does not follow test procedure correctly | Hipot test is not executed correctly or unsafely | N/A | 9 | Operator doesn't follow procedure - cables not connected per setup | 10 | Operator trained and signed off on procedure per G-QAP017. Test instrument is automated with Pass/Fail result and test will fail if the hipot or ground cable is not connected | 1 | 90 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 9 | 5 | 1 | 45
|  |  |  | Test Equipment: Hipot Tester | Test Equipment: Hipot Tester does not test unit properly or has the incorrect voltage/trip current settings | Hipot test is not executed correctly | N/A | 9 | Hipot Tester malfunction or out of calibration | 4 | Process is validated - LR0005; Equipment is on a calibration schedule. Hipot tester settings are checked on a daily basis before use and recorded on log 10060262. | 1 | 36 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Hipot Tester cables | Test Equipment: Hipot Tester cables are open circuit | Hipot test is not executed correctly - potential shock hazard | N/A | 9 | broken wire in cable | 10 | Cables are checked on a daily basis before use and test will fail if the hipot or ground cable is not connected or broken. Preventive Maintenance G-AMFG-1180-00 inspects Hipot cable assy | 1 | 90 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 9 | 4 | 1 | 36
|  | Leakage Current Tests | Leakage Current Tests | PB500 ELECTRICAL SAFETY TEST & BUZZER CARD CHARGE PROCEDURE - 10076002 | Procedures misinterpreted due lack of clarity; incorrectly written | Leakage Current test is not executed correctly or unsafely | N/A | 6 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow test procedure correctly | Leakage Current tests are not executed correctly | N/A | 6 | Operator doesn't follow procedure - does not connect cables per setup | 10 | Operator trained and signed off on procedure per G-QAP017. Test is automated, but needs cables connected properly. Hipot test should catch any gross enclosure leakage faults which would be seen at the USB connectors | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment:Safety Analyzer Tester | Test Equipment: Safety Analyzer Tester does not test unit properly or has the incorrect settings | Leakage Current tests are not executed correctly | N/A | 6 | Safety Analyzer Tester malfunction or out of calibration. | 10 | Process is validated - LR0005; Equipment is on a calibration schedule. Settings are stored in the tester unit and recalled each time the unit is turned on. | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Safety Analyzer Tester cables | Test Equipment: Safety Analyzer Tester cables are open circuit | Leakage Current tests are not executed correctly | N/A | 6 | Broken wire in cable | 10 | Cables are checked for continuity on a daily basis before use | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  | Software Download incl Battery Patch Download | Software Download incl Battery Patch Download | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Software is not downloaded or verified to be downloaded correctly | N/A | 6 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; process is validated - LR0005; Successful download message is checked when download is complete. Test process will not execute successfully without product software loaded and S/W version are recorded on DHR later in procedure | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow test procedure | Operator does not complete product software download successfully | N/A | 6 | Operator doesn't follow procedure - power switch in wrong position for ST10 download, wrong com port selected for either download app, download cable not connected | 10 | Operator trained and signed off on procedure per G-QAP017. Process is validated - LR0005; Successful download message is checked when download is complete. Download app cannot start with wrong com port number. Test process will not execute successfully without product software loaded | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Download cable, USB Type A/ Mini B | Cable: download cable shorted or open circuit | Software download does not execute | N/A | 4 | Cable is worn or broken from use | 2 | Process is validated - LR0005; Download cannot start with bad cable. Test process will not run without product software loaded. Download cable can be replaced with off the self spare. Preventive Maintenance G-AMFG-1180-00 inspects download cable | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Download PC | Test Equipment: Download PC does not function properly - download app does not run | Software download does not execute | N/A | 4 | PC software is corrupt or PC hardware problem | 2 | Process is validated - LR0005; Successful download message when download is complete and S/W version is recorded on DHR later in procedure. Test process will not run without product software loaded. PC can be replaced with any spare. | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Product Software CPU: p/n 10039034 Power: p/n 3832100 | Product Software is not the correct version | Vent has incorrect version of software loaded | N/A | 4 | Download PC has relevant software in incorrect file location. | 2 | Process is validated - LR020. Product software is ECO controlled/released. Product s/w version is recorded on DHR later in procedure. | 2 | 16 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | PB540/PB520 CPU S/W is loaded instead of the PB560S/W | Vent has incorrect software loaded | N/A | 4 | Incorrect file path is chosen at download. | 3 | CPU software version is hand written on DHR. Operator must record downloaded version and verify against the version printed on the DHR | 2 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Incorrect CPU software is loaded on vent | vent has incorrect software | N/A | 4 | Incorrect file path is chosen at download. | 3 | Openvent software at final test will look for correct PB560 vent serial number. | 2 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Incorrect version Power Management Board S/W is loaded | Generic Power management software for PB500 units | N/A | 1 | Universal PIC, Power management software | 1 | Process validated as per LR0017. PIC software (power management) will be recorded and verified by operator. | 1 | 1 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Download Software p/n 3832500 and 3832600 | Download Software performs download incorrectly | Software download fails | N/A | 4 | unexpected software fault | 2 | Process is validated - LR0005; Successful download message is checked when download is complete. Product S/W version is recorded on DHR later in procedure. Test process will not execute successfully without product software loaded | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Download Software p/n 3832500 and 3832600 | Download Software is not the correct version | Software download may not work as expected (specific to reason for update) | N/A | 4 | specific to reason for update | 2 | Process is validated - LR0005. Download software is ECO controlled/realeased. A validation process is performed with update of download software. Product S/W version is recorded on DHR later in procedure. | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Battery Patch Software is not downloaded per procedure | Battery Capacity Check will fail if the Battery Patch s/w download was required. | N/A | 6 | Operator does not complete download successfully | 4 | Operator trained and signed off on procedure per G-QAP017. If not caught at the Battery Capacity Check, may also fail the Battery Discharge-Charge Cycle Testing | 2 | 48 | N/A | N/A | N/A | - | - | - | -
|  | Check Power Up Indicators | Check Power Up Indicators | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Power Up indicators are not checked correctly | N/A | 6 | Lack of review of procedure | 10 | The AC Led is also checked in Power Source Switching test, the Yellow alarm led is also checked in the Remote Alarm test and the Red alarm led is also checked in the Patient Disconnect Test of the 10036007 test procedure. The two buzzers are also checked in the 10036007 test proc. Procedure is ECO controlled in Agile; Process is validated - LR0005; operator trained and signed off on procedure | 2 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 6 | 2 | 2 | 24
|  |  |  | Operator | Operator: does not follow test procedure | Power Up indicators are not checked to test procedure | N/A | 6 | Operator doesn't follow procedure - operator does not verify all indicators | 10 | Operator trained and signed off on procedure per G-QAP017. The AC Led and is also checked in Power Source Switching test, the yellow alarm led is also checked in the Remote Alarm test and the Red alarm led is also checked in the Patient Disconnect Test of the 10036007 test procedure. The two buzzers are also checked in the 10036007 test proc. | 2 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 6 | 6 | 2 | 72
|  | Set Vent S/N | Set Vent S/N | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Vent has wrong or no S/N stored | N/A | 2 | Lack of review of procedure | 10 | S/N written is verified later in procedure, Procedure is ECO controlled in Agile; Process is validated - LR0004/LR0005; operator trained and signed off on procedure | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not enter S/N correctly | Vent has wrong S/N stored | N/A | 2 | Operator doesn't follow procedure - enters wrong S/N or does not click Ecruiture S/N to store serial number | 10 | S/N written is verified later in procedure, operator trained and signed off on procedure per G-QAP017. | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  | Setup Time/Date | Setup Time/Date | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Time/date not written to ventilator | N/A | 5 | Lack of review of procedure | 10 | Time/date written verified later in procedure, Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 100 | Review PFMEA as part of G-QAP056. Review procedure vs practice. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 5 | 2 | 2 | 20
|  |  |  | Operator | Operator: does not enter time or date correctly | Wrong time or date entered to vent | N/A | 5 | Operator doesn't follow procedure - wrong time and/or date entered | 10 | Time/date is written verified later in procedure, operator trained and signed off on procedure per G-QAP017. | 2 | 100 | Review PFMEA as part of G-QAP056. Review procedure vs practice. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 5 | 5 | 2 | 50
|  | Record Software Versions | Record Software Versions | Operator | Operator: does not record software version numbers as displayed on ventilator screen | Test not executed correctly - incorrect software s/n's recorded | N/A | 1 | Operator doesn't follow procedure - wrong number recorded - typo errors | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; operator trained and signed off on procedure per G-QAP017. | 2 | 20 | N/A | N/A | N/A | - | - | - | -
|  | Check Vent S/N | Check Vent S/N | Operator | Operator: does not check that the ventilator S/N stored matches the DHR correctly | Test not executed correctly - potential for missed faults | N/A | 1 | Operator doesn't follow procedure - comparison made incorrectly | 10 | Operator trained and signed off on procedure per G-QAP017. Checked in Product Authorisation Release | 2 | 20 | N/A | N/A | N/A | - | - | - | -
|  | Remote Alarm Tests | Remote Alarm Tests | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Remote Alarm is not checked correctly | N/A | 6 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow procedure correctly | Remote Alarm is not checked to test procedure correctly | N/A | 6 | Operator doesn't follow procedure - does not verify all indicators or the DVM continuity checks | 10 | The normally closed and open Remote Alarm Contacts are checked using a DVM for continuity in both the deactivated and activated positions and recorded on DHR. Operator trained and signed off on procedure per G-QAP017. | 2 | 120 | Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 6 | 2 | 2 | 24
|  |  |  | Test Equipment: Patient Connector Assembly #1 | Patient Connector Assembly #1 hoses leaking air | Ventilator producing unexpected alarms causing confusion when checking expected alarms | N/A | 1 | Patient Connector Assembly #1 has broken air hoses or fittings | 5 | Process is validated - LR0005; Preventive Maintenance Procedure G-AMFG-1180-00 inspects tubing/fitting | 3 | 15 | N/A | N/A | N/A | - | - | - | -
CTQ |  | Power Source Switching Tests | Power Source Switching Tests | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Power Source switching is not checked correctly | N/A | 6 | Lack of review of procedure | 2 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow test procedure | Power Source switching is not checked to test procedure correctly | N/A | 6 | Operator doesn't follow procedure - setup incorrect or makes power supply adjustment-measurements incorrectly | 10 | Operator trained and signed off on procedure per G-QAP017. DHR 10038461 includes all checks performed in procedure. Measurements checked against limits in Product Authorisation Release. Power Management board has passed similar tests for power switching in functional test | 2 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 6 | 2 | 2 | 24
|  |  |  | Operator | Operator: the power supply is adjusted too fast | External DC Power Supply reading recorded is not the correct value or there is not uniformity of testing between test operators | N/A | 6 | Operator doesn't follow procedure - makes power supply adjustment too fast | 10 | Test procedure contains steps to increment/decrement power supply by 0.1V steps, then stop and observe for power led change. Operator trained and signed off on procedure per G-QAP017. Measurements checked against limits in Product Authorisation Release. Power Management board has passed similar tests for power switching in functional test | 2 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 6 | 4 | 2 | 48
|  |  |  | Test Equipment: External DC Adjustable Power Supply | External DC Adjustable Power Supply output is out of spec, or not displaying voltage accurately | Power Source switching is not checked to test procedure correctly - false test failures or missed faults | N/A | 6 | External DC Adjustable Power Supply out of calibration or faulty | 4 | Process is validated - LR0005; Equipment is on a calibration schedule. Spare PS available | 3 | 72 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Test DC Power Cable, P/N 10022945 | Test DC Power Cable, P/N 10022945 is broken | Power Source switching is not checked to test procedure correctly - false test failures | N/A | 5 | Broken wire in cable | 4 | Preventive Maintenance Procedure G-AMFG-1180-00 inspects DC power cable. Process is validated - LR0005; | 2 | 40 | N/A | N/A | N/A | - | - | - | -
CTQ |  | Alarm tests | Alarm tests | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Alarm tests not executed correctly | N/A | 7 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0004/LR0005; | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 2 | 2 | 28
|  |  |  | Operator | Operator: does not follow test procedure | Alarm tests not executed to test procedure correctly - potential for missed faults | N/A | 7 | Operator doesn't follow procedure - checks for leds, alarm sound and duration not verified correctly | 10 | DHR 10038461 includes checks to be performed. Operator trained and signed off on procedure per G-QAP017. Measurement checked against limits in Product Authorisation Release. | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 5 | 2 | 70
|  |  |  | Test Equipment: Sound Level Meter | The Alarm sound level is reading incorrectly | Alarm tests not executed correctly - potential for missed faults or false failues | N/A | 7 | Sound level meter is out of calibration | 1 | Process is validated - LR0004/LR0005; Equipment is on a calibration schedule, Sound Level Meter is checked on a daily basis and recorded on log 10039028 | 2 | 14 | N/A | N/A | N/A | - | - | - | -
|  |  |  | SPKR FAULT1 or SPKR FAULT4 alarm events | The buzzer battery is not sufficently charged for the test | False failure for the Involuntary Alarm Test | N/A | 6 | The buzzer battery is not sufficently charged for the test | 4 | Procedure is ECO controlled in Agile; Process is validated - LR0004/LR0005; 10036007 test proc has test note to insure vent is switched on for at least 15 minutest to permit ample charging of the buzzer battery before performing the Involuntary Stop Alarm Test | 3 | 72 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Exhalation Valve Alarm does not sound as expected in test | Exhalation Valve Alarm test is not executed correctly | N/A | 7 | PB560 Vent settings are not set correctly to be in a valve type configuration per procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0004/LR0005. Vent settings are specifically stated in procedure to insure vent will inidcate an Exhalation Valve Alarm | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | No non value type configuration used in this process. Reviewed procedure vs. practice during validation build. No occurrence detected during the build. | 7 | 4 | 2 | 56
|  | USB Ports Test | USB Ports Test | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | USB not functioning | N/A | 6 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | No impact on ventilation funciton. USB port used only for downloading logs. Reviewed practice vs. procedure during validation build. No occurrence detected during the build. | 4 | 5 | 2 | 40
|  |  |  | Operator | Operator: does not follow test procedure correctly. USB Memory Device is not inserted into both ports | USB not functioning | N/A | 6 | Operator doesn't follow procedure - both ports are not tested | 10 | Operator trained and signed off on procedure per G-QAP017. 10038461 DHR has check boxes to indicate each port is tested | 2 | 120 | Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | No impact on ventilation funciton. USB port used only for downloading logs. Reviewed practice vs. procedure during validation build. No occurrence detected during the build. | 4 | 5 | 2 | 40
|  |  |  | Test Equipment - USB Memory Device | USB Device has an fault or is not formatted to the FAT32 format | Vent does not recognise the USB memory device | N/A | 6 | USB Device has an fault or is not formatted the the FAT32 format | 2 | Process is validated - LR0005; USB Memory device can be easily checked on any PC | 3 | 36 | N/A | N/A | N/A | - | - | - | -
|  | Calibrate Patient Sensor | Calibrate Patient Sensor | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Patient Pressure sensor not calibrated correctly | N/A | 7 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 2 | 2 | 28
|  |  |  | Operator | Operator: does not follow test procedure correctly | Patient Pressure sensor not calibrated correctly -masurements are not accurate | N/A | 7 | Operator doesn't follow procedure - pressure is not set to 40cmH20 for calibration. PTS2000 is powered on with pressure applied to input | 10 | Operator trained and signed off on procedure per G-QAP017. Presurre is checked immediatley after cal and verified later in Final Test Procedure. Measurement checked against limits in Product Authorisation Release. Procedure contains warning to not apply pressure during power-up. PTS2000 is not power cycled during any tesing. | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 2 | 2 | 28
|  |  |  | Test Equipment: Patient Connector Assembly #2 | Test Equipment: Air tubing leaking air | Patient Pressure sensor not calibrated correctly -masurements are not accurate | N/A | 7 | Broken or cracked air hose | 3 | Preventive Maintenance Procedure G-AMFG-1180-00 inspects tubing/fitting. Process is validated - LR0005; Pressure is checked in the Final Test Procedure using different air tubing | 3 | 63 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: PTS2000 | Test Equipment: PTS2000 is not reading accurately | Patient Pressure sensor not calibrated correctly-measurements are not accurate | N/A | 7 | PTS2000 is out of calibration | 1 | Process is validated - LR0004; Equipment is on a 3 month internal calibration schedule and 12 month ext calibration . | 3 | 21 | N/A | N/A | N/A | - | - | - | -
|  | Calibrate Exhalation Valve Sensor | Calibrate Exhalation Valve Sensor | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Exhalation valve sensor not calibrated correctly | N/A | 7 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; Basic Exhalation Valve opereation is checked in Final System Test | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 2 | 2 | 28
|  |  |  | Operator | Operator: does not follow test procedure & updated connector assembly #2 is not fitted correctly as per procedure 10036007 rev G | Exhalation valve sensor not calibrated correctly. Measurements are not accurate | N/A | 7 | Operator doesn't follow procedure - fails without detection by missing the long beep response after calibration | 10 | Operator trained and signed off on procedure per G-QAP017; Exhalation Valve opereation is checked in Final System Test. Simple process with little possibility for operator mistakes. Operators trained on procedure 10036007 rev G | 2 | 140 | Update procedure & DHR to add a check to ensure the connector #2 assembly is fitted correctly. | Dara O'Connor | Update procedure & DHR on ECO-R197243 | 7 | 10 | 1 | 70
|  | Calibrate Inspiratory Flow Sensor | Calibrate Inspiratory Flow Sensor | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Inspiratory Flow sensor not calibrated correctly | N/A | 7 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; Flow rates are verified in this test section | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 2 | 2 | 28
|  |  |  | Operator | Operator: does not follow test procedure | Inspiratory Flow sensor not calibrated correctly | N/A | 7 | Operator doesn't follow procedure - calibration not accurate or record wrong flow reading, typos | 10 | Operator trained and signed off on procedure per G-QAP017 . Flow rates are checked after calibration in this section., and volume tests are done in Final Test. Measurements checked against limits in Product Authorisation Release. | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 4 | 2 | 56
|  |  |  | Test Equipment: 22mm air tubing | Test Equipment: Air tubing leaking air | Inspiratory Flow sensor not calibrated correctly-measurements are not accurate | N/A | 7 | Broken or cracked air hose | 3 | Preventive Maintenance Procedure G-AMFG-1180-00 inspects air tubing. Process is validated - LR0005; | 3 | 63 | N/A | N/A | N/A | - | - | - | -
|  | Turbine Performance Tests | Turbine Performance Tests | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Turbine performance is not checked properly | N/A | 7 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0004; | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 2 | 2 | 28
|  |  |  | Operator: | Operator: does not follow test procedure correctly | Turbine performance is not checked properly - potential for missed faults or false failures or inaccurate measurements | N/A | 7 | Operator doesn't follow procedure - setup, mis-reading PTS2000, recording incorrect value. PTS2000 is powered on with pressure applied to input | 10 | Operator trained and signed off on procedure per G-QAP017. Measurements checked against limits in Product Authorisation Release. Turbine failures would be caught in Final System Test in the Volume accuracy section (1500mL). Procedure contains warning to not apply pressure during power-up. PTS2000 is not power cycled during any tesing. | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 5 | 2 | 70
|  |  |  | Test Equipment: Patient Connector Assembly #2 | Test Equipment: Air tubing leaking air | Turbine performance is not checked properly-measurements are not accurate | N/A | 7 | Broken or cracked air hose | 4 | Preventive Maintenance Procedure G-AMFG-1180-00 inspects air tubing/fitting. Process is validated - LR0005; | 2 | 56 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: PTS2000 | Test Equipment: PTS2000 is not reading accurately | Turbine performance is not checked properly-measurements are not accurate | N/A | 7 | PTS2000 is out of calibration | 3 | Process is validated - LR0005; Equipment is on a 3 month internal calibration schedule and 12 month ext calibration . | 3 | 63 | N/A | N/A | N/A | - | - | - | -
|  | Battery Discharge-Charge Cycle | Battery Discharge-Charge Cycle | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Battery discharge-charge cycle not run correctly with missing low battery/empty battery alarms or battery not fully charging | N/A | 5 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 100 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 5 | 2 | 2 | 20
|  |  |  | Operator: | Operator: does not follow test procedure | Battery discharge-charge cycle not run correctly with missing low battery/empty battery alarms or battery not fully charging | N/A | 5 | Operator doesn't follow procedure - inccorrect vent settings set, or not verifying log time between low and empty battery alarms correctly | 10 | Low and Empty Battery Alarm Event times are recorded on DHR and checked by Product release Authorisation. Operator trained and signed off on procedure per G-QAP017. Alarm time and % charged are recorded and checked by Product release Authorisation Authorisation | 3 | 150 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 5 | 5 | 3 | 75
|  |  |  | Test Equipment: AC Timer | Test Equipment: incorrect discharge or charge times | Battery discharge-charge cycle not run correctly with missing low battery/empty battery alarms or battery not fully charging | N/A | 5 | Inproperly set or faulty AC timer | 3 | Timer cover prevents tampering. 10036007 procedure includes Daily Check for timer setting check. Process is validated - LR0005; | 4 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Patient Connector Assembly #1 | Test Equipment: Air tubing leaking air | Unexpected alarms due to air leaks | N/A | 5 | Broken or cracked air hose | 3 | Preventive Maintenance Procedure G-AMFG-1180-00 inspects tubing/fitting. Process is validated - LR0005; | 4 | 60 | N/A | N/A | N/A | - | - | - | -
|  | System Measurements Check | System Measurements Check | Operator | Operator: reads/records the wrong value or does not make measurements immediately after the overnight battery discharge/charge cycle | Test not executed correctly - potential for missed faults or false failures | N/A | 6 | Operator doesn't follow procedure - reads measurement incorrectly or typo errors in recording | 4 | Operator trained and signed off on procedure per G-QAP017. Measurements are checked against limits in Product Authorisation Release | 3 | 72 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Incorrect limits specified | N/A | 6 | Lack of review of procedure | 2 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 24 | N/A | N/A | N/A | - | - | - | -
CTQ |  | Check Battery | Check Battery | Operator | Operator: reads/records the wrong value or does not make measurements immediately after the overnight battery discharge/charge cycle | Battery information is not checked correctly | N/A | 7 | Operator doesn't follow procedure - operator mis-reads or records bad data as good data | 10 | Operator trained and signed off on procedure. Measurements recorded on DHR and checked by Product release Authorisation Authorisation | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 4 | 2 | 56
|  |  |  | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Battery information is not checked properly-incorrect limits used | N/A | 7 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 2 | 2 | 28
CTQ |  | Alarm (buzzer) sound level tests | Alarm (buzzer) sound level tests | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Alarm sound level test is not executed correctly | N/A | 8 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow test procedure | Alarm sound level test not executed to test procedure correctly - potential for missed faults or false failures | N/A | 8 | Operator doesn't follow procedure - operator misreads a bad Sound Level Meter measurement as good, or has the sound level meter in the wrong position - false failure, or does not have the sound level meter Range set per procedure | 10 | Operator trained and signed off on procedure which clearly states where to position the sound level meter. Sound level has easy to read digital display | 1 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Sound Level Meter | The Alarm sound level is reading incorrectly | Alarm sound level - potential for missed faults or false failues | N/A | 8 | Sound level meter is out of calibration | 1 | Process is validated - LR0005; Equipment is on a calibration schedule. Sound Level Meter is checked on a daily basis and recorded on log 10039028 | 2 | 16 | N/A | N/A | N/A | - | - | - | -
|  | Faults Log Check | Faults Log Check | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Faults Log is not checked properly - list of applicable faults is not complete | N/A | 5 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 100 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 5 | 2 | 2 | 20
|  |  |  | Operator | Operator: does not follow test procedure correctly | Important fault is missed so vent is failed for normal operating fault | N/A | 5 | Operator doesn't follow procedure - operator reads/records information incorrectly or missses importamt log events | 10 | Operator trained and signed off on procedure per G-QAP017. Fault log is checked again in Final System test procedure 10039030 | 3 | 150 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 5 | 5 | 3 | 150
|  | Battery Discharge-Charge Cycle Event Times | Battery Discharge-Charge Cycle Event Times | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Alarm Event times are not recorded and/or evaluated properly | N/A | 8 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 8 | 2 | 2 | 32
|  |  |  | Operator | Operator: does not follow test procedure; math error | Alarm Event times are not recorded and/or evaluated properly (math mistake) | N/A | 8 | Operator does not record the corrrect event time or makes a math mistake calculating the low to involuntay stop or empty to involuntary stop alarm events | 10 | Operator trained and signed off on procedure per G-QAP017. Fault log is checked again in Final System test. Measurements recorded on DHR and checked by Product release Authorisation Authorisation | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 8 | 5 | 2 | 80
|  |  |  | PB PVTS Downlaod s/w Tool p/n 100322788 | Download Software is not the correct version | Alarm Event times may not be accurately displayed for printout | N/A | 8 | specific to reason for update | 2 | Process is validated - LR0005. PB PVTS software is validated and ECO controlled/realeased, A validation process is performed with update of PVTS software. | 3 | 48 | N/A | N/A | N/A | - | - | - | -
|  | Vent Preliminary Checks | Vent Preliminary Checks | Procedure: PB560, System Final Test - P/N 10039030 | Procedures misinterpreted due lack of clarity; incorrectly written | Paperwork or vent exterior checks are performed correctly | N/A | 4 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 80 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow test procedure correctly | Paperwork or vent exterior checks are not performed correctly | N/A | 4 | Operator doesn't follow procedure - checks not made thoroughly | 10 | Operator trained and signed off on procedure per G-QAP017. S/n checked at Pack process, Calibration/DHR checked by Product release Authorisation, Screw check in this step is an addional check to assembly | 2 | 80 | N/A | N/A | N/A | - | - | - | -
|  | Real Time Clock Tests | Real Time Clock Tests | Procedure: PB560, System Final Test - P/N 10039030 | Procedures misinterpreted due lack of clarity; incorrectly written | Time and Date are not verified accurately | N/A | 3 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 4 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 3 | 2 | 4 | 24
|  |  |  | Operator: | Operator: does not follow test procedure correctly | Time and Date are not verified accurately | N/A | 3 | Operator doesn't follow procedure in misreading time or date | 10 | Operator trained and signed off on procedure. Simple process to read date/time | 4 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 3 | 4 | 4 | 48
CTQ |  | O2 Valve Leak Tests | O2 Valve Leak Tests | Procedure: PB560, System Final Test - P/N 10039030 | Procedures misinterpreted due lack of clarity; incorrectly written | O2 Valve Leak test is not executed correctly | N/A | 7 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 2 | 2 | 28
|  |  |  | Operator | Operator: does not follow test procedure correctly | O2 Valve Leak test is not executed correctly | N/A | 7 | Operator doesn't follow procedure - setup | 10 | Operator trained and signed off on procedure. Test is automated. | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 4 | 2 | 56
|  |  |  | Test Equipment: Leak Tester | Test Equipment: Leak Tester does not function as expected | O2 Valve Leak test is not executed correctly | N/A | 7 | Leak Tester has the incorrect settings, tester fault or calibration error | 2 | Process is validated - LR0005; Equipment is on a calibration schedule. Settings are stored in the tester unit. 10023391 O2 Valve Leak Test doc describes setup. Spare Leak Tester available | 4 | 56 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Leak Tester output air tubing/couplings | Test Equipment: Leak Tester output air tubing/couplings leaks air | O2 Valve Leak test is not executed correctly - false failures | N/A | 4 | Test Equipment: Leak Tester output air tubing/couplings is damaged | 2 | Process is validated - LR0005; leak in output tubing/fitting will fail 'Fill' poortion of automated test routine. Slower rate leaks will fail test repeatdly. Preventive Maintenance Procedure G-AMFG-1180-00 inspects air tubing/couplings | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Air Supply | Air Supply: Pressure not correct | O2 Valve Leak test is not executed correctly if air pressure is low or leak tester may be damaged is air pressure is too high | N/A | 7 | Air Supply: Pressure not between 110-140psi | 4 | Process is validated - LR0005; Air pressure is sourced from a regulated supply. Daily check verifies air supply pressure and recorded on log p/n 10039034. Valve Leak Test doc p/n 100233912 describes setup. | 2 | 56 | N/A | N/A | N/A | - | - | - | -
CTQ |  | Volume and Pressure Tests: Calibrate Expiratory Flow Sensor | Volume and Pressure Tests: Calibrate Expiratory Flow Sensor | Procedure: PB560, Calibration and Test - P/N 10036007 | Procedures misinterpreted due lack of clarity; incorrectly written | Expiratory Flow sensor not calibrated correctly | N/A | 7 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; Flow rates are checked in the section | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 7 | 4 | 2 | 56
|  |  |  | Operator | Operator: does not follow test procedure | Expiratory Flow sensor not calibrated correctly | N/A | 7 | Operator doesn't follow procedure - missing a CAL FAULT message on vent | 10 | Operator trained and signed off on procedure per G-QAP017 . Simple process with little possibility for operator mistakes. Expiratory volume measurement checks calibration | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed practice vs. procedure during validation build. No occurrence detected during the build. | 7 | 4 | 2 | 56
|  |  |  | Operator | Operator: does not select Pediatric Patient Circuit to Yes option in the vent preferences menu for calibrating and testing using the Pediatric patient circuit | Expiratory Flow sensor not calibrated correctly for Pediatric, pediatric volume/pressure tests may fail | N/A | 7 | Operator does not follow procedure - ingores step | 10 | Operator trained and signed off on procedure per G-QAP017 . Check box added in DHR 10038461 for selecting the Pediatric Patient Circuit option in the vent preferences menu | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed practice vs. procedure during validation build. No occurrence detected during the build. | 7 | 4 | 2 | 56
|  |  |  | Operator | Operator: does not select Pediatric Patient Circuit to No in the vent preferences menu after Pediatric testing | Expiratory Flow sensor not calibrated correctly for Adult, adult volume/pressure tests may fail | N/A | 7 | Operator does not follow procedure - ingores step | 10 | Operator trained and signed off on procedure per G-QAP017 . Check box added in DHR 10038461 for de-selecting the Pediatric Patient Circuit option in the vent preferences menu | 2 | 140 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed practice vs. procedure during validation build. No occurrence detected during the build. | 7 | 4 | 2 | 56
|  |  |  | Test Equipment: Patient Circuit | Test Equipment: Air tubing leaking air | Expiratory Flow sensor not calibrated correctly-measurements are not accurate | N/A | 7 | Broken or cracked air hose | 4 | Test procedure 10039030 has a daily inspection of tubing. Process is validated - LR0005 | 2 | 56 | N/A | N/A | N/A | - | - | - | -
CTQ |  | Volume and Pressure Tests | Volume and Pressure Tests | Procedure: PB560, System Final Test - P/N 10039030 | Procedures misinterpreted due lack of clarity; incorrectly written | Volume or Pressure Test is not executed correctly | N/A | 8 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0004/LR0005; | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 8 | 3 | 2 | 48
|  |  |  | Open Vent Test Software Tool p/n 10026042 | Incorrect version of Open Vent S/W used | Volume or Pressure Test is not executed correctly | N/A | 8 | Lack of SW version control | 3 | Software is ECO controlled in Agile; Process is validated - LR0004/LR0005; | 3 | 72 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Open Vent Test Software Tool p/n 10026042 | s/w tool causes test a blank test report sheet to be produced | Volume or Pressure Test is not executed completely | N/A | 5 | Lack of s/w tool testing | 5 | Software has error handling code so that if there are any com errors, a test report will not be produced | 2 | 50 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Software: Ped Test Script P/N 10039032 Adult Test Script P/N 10039033 | Incorrect version of Test Script used | Volume or Pressure Test is not executed correctly | N/A | 8 | Lack of document (script) version control | 3 | Scripts are ECO controlled in Agile; Process is validated - LR0004/LR0005; | 3 | 72 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Software: Ped Test Script P/N 10039032 Adult Test Script P/N 10039033 | Test Script has test command written wrong | Volume and/or Pressure Test is not executed correctly | N/A | 8 | Lack of review of test script | 2 | Test Script is peer reviewed and controlled in Agile; Process is validated - LR0004/LR0005; | 3 | 48 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator: | Pediatric and Adult Tests: Proximal Pressure tubing from vent to PTS2000 low pressure input is not connected for test | Pressure test is not executed correctly - false failures will be reported for pressure tests in both Ped and Adult scripts. | N/A | 2 | Operator doesn't follow procedure to connect pressure tubing to PTS2000 | 10 | Operator trained and signed off on procedure per G-QAP017. Vent will display 'Check Prox Line' error and test report will fail pressure tests. Pediatric test script will additionally fail the Safety Pressure Test. | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator: | Pediatric Tests: The Adult test lung is connected instead of the Pediatric Test lung | Volume and Pressure tests are not executed correctly - false failures will be reported for volume and pressure tests | N/A | 2 | Operator doesn't follow procedure to connect Adult test lung | 10 | Operator trained and signed off on procedure per G-QAP017. Vent will display 'Low Pressure Disconnect ' alarm and test report will indicate a Disconnection Alarm and Safety Pressure failures | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator: | Pediatric Tests: The Adult test script is executed instead of the Pediatric Test script | Volume and Pressure tests are not executed correctly - false failures will be reported for volume and pressure tests | N/A | 2 | Operator doesn't follow procedure to execute pediatric test script as directed | 10 | Operator trained and signed off on procedure per G-QAP017. The Vent will indicate a 'High Pressure ' alarm and the test report will indicate a 'High Pressure' failure | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator: | Pediatric and Adult Tests: The exhl valve control pressure line is not connected to vent insp block | Volume and Pressure tests are not executed correctly - false failures will be reported for volume and pressure tests | N/A | 2 | Operator doesn't follow procedure to connect the exhl valve as directed | 10 | Operator trained and signed off on procedure per G-QAP017. The Vent will indicate a 'Check Exhl Valve' alarm. | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator: | Pediatric and Adult Tests: The test lung is not connected | Volume and Pressure tests are not executed correctly - false failures will be reported for volume and pressure tests | N/A | 2 | Operator doesn't follow procedure to connect the appropiate test lung as directed | 10 | Operator trained and signed off on procedure per G-QAP017. The Vent will indicate a 'Low Pressure ' alarm and the test report will indicate a 'Disconnection Alarm' failure | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator: | Adult Tests: The pediatric test lung is connected instead of the adult test lung connected | Volume and Pressure tests are not executed correctly - false failures will be reported for volume and pressure tests | N/A | 2 | Operator doesn't follow procedure to connect the adult test lung as directed | 10 | Operator trained and signed off on procedure per G-QAP017. The Vent will indicate a 'High Pressure ' alarm and the test report will indicate a 'High Pressure' failure in volume tests | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator: | Adult Tests: The pediatric test script is executed instead of the adult test script | Volume and Pressure tests are not executed correctly - false failures will be reported for volume and pressure tests | N/A | 2 | Operator doesn't follow procedure to execute the adult test script as directed | 10 | Operator trained and signed off on procedure per G-QAP017. The Vent will indicate a 'Low Pressure ' alarm and the test report will indicate a 'Disconnection Alarm' failure | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator: | A Fail Test result on the test report is missed when assessing the overall test result | False Pass | N/A | 8 | Operator misses a Fail Result | 10 | Operator trained and signed off on procedure per G-QAP017. Measurements recorded on DHR and checked by Product release Authorisation. | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 8 | 2 | 2 | 32
|  |  |  | Test Equipment: PTS2000 | Test Equipment: PTS2000 does not measure pressure accurately | Pressure Control Test is not executed correctly - Pressure Mode of vent is not checked to test procedure | N/A | 8 | PTS2000 Null Pressure is not at zero | 3 | Pediatric and Adult Test Scripts performs a zero PTS200 pressure command at the beginning of each test script. | 2 | 48 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: PTS2000 | Test Equipment: PTS2000 does not measure pressure or volume accurately | Volume and Pressure tests are not executed correctly - false failures will be reported for volume and pressure tests | N/A | 8 | PTS2000 is out of calibration or equipment fault. | 1 | Process is validated - LR0004/LR0005; Equipment is on a 3 month internal calibration schedule and 12 month ext calibration . | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: Patient Circuit and air tubing/couplings | Test Equipment: Patient Circuit and air tubing/couplings leaks air | Pressure or Volume Test is not executed correctly - false failures or false pass | N/A | 9 | Test Equipment: Patient Circuit and air tubing/couplings is damaged | 2 | Process is validated - LR0004/LR0005; Patient tubing and exhl valve is on a 2 week replacement schedule. Preventive Maintenance Procedure G-AMFG-1180-00 inspects air tubing/fittings. | 3 | 54 | N/A | N/A | N/A | - | - | - | -
|  | FiO2 Sensor Detection Test | FiO2 Sensor Detection Test | Procedure: PB560, System Final Test - P/N 10039030 | Procedures misinterpreted due lack of clarity; incorrectly written | FiO2 Sensor Detection Test function is not checked correctly | N/A | 8 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed procedure vs practice during validation build. No occurence detected during build. | 8 | 4 | 2 | 48
|  |  |  | Operator | FiO2 Sensor Detection Test function is not checked correctly | FiO2 Sensor Detection is incorrectly passed. | N/A | 8 | Operator doesn't follow procedure - FiO2 sensor/cable setup . | 10 | Operator trained and signed off on procedure per G-QAP017 | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed procedure vs practice during validation build. No occurence detected during build. | 8 | 4 | 2 | 48
|  |  |  | Operator | FiO2 Sensor 'use by' date is not checked. | FiO2 Sensor Detection is incorrectly passed. | N/A | 8 | FiO2 Sensor is faulty or past service life | 10 | Spare (Airox) FiO2 Sensor is available. FiO2 Sensor is marked with end of use label. Length of test use is reduced from recommended Service Life term | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | FI02 Sensor is added to the shelf life procedure G-AMFG-1066-00. Reviewed practice vs. procedure during validation build. No occurrence detected during build. | 8 | 4 | 2 | 48
|  | O2 Valve Funtional Tests | O2 Valve Funtional Tests | Procedure: PB560, System Final Test - P/N 10039030 | Procedures misinterpreted due lack of clarity; incorrectly written | O2 Valve Functional test is not executed correctly | N/A | 8 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0004; | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 8 | 2 | 2 | 32
|  |  |  | Operator | Operator: does not follow test procedure correctly | O2 Valve Functional test is not executed correctly according to test procedure. | N/A | 8 | Operator doesn't follow procedure - setup, misreads a bad PTS2000 value as good or records wrong values. | 10 | Operator trained and signed off on procedure per G-QAP017. Measurements recorded on DHR and checked by Product release Authorisation. | 2 | 160 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. | PB560 Manufacturing Engineer. | Reviewed historical failure yield data and PFMEA for PB540. Based on this data, reduce occurrence rate in line with PB540 | 8 | 3 | 2 | 48
|  |  |  | Test Equipment: PTS2000 | Test Equipment: PTS2000 does not measure % O2 accurately | Pressure Control Test is not executed correctly - O2 valve is not checked to test procedure | N/A | 8 | PTS2000 is out of calibration or equipment fault | 1 | Process is validated - LR0005; Equipment is on a 3 month internal calibration schedule and 12 month ext calibration . | 3 | 24 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Test Equipment: O2 tubing/couplings | Test Equipment: O2 tubing/couplings leaks O2 gas | O2 Valve Functional test is not executed correctly - false failures | N/A | 5 | O2 tubing/couplings is damaged | 4 | Process is validated - LR0005; O2 leak would cause a false failure. Preventive Maintenance Procedure G-AMFG-1180-00 inspects air tubing/fitting | 3 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | O2 Supply | O2 Supply: Pressure not correct | O2 Valve Functional test is not executed correctly if O2 pressure is low | N/A | 8 | O2 supply is turned off or pressure is set to low | 4 | Process is validated - LR0005; O2 pressure is sourced from a regulated supply. Daily check verifies O2 supply setting and reocrded on log 10039304 | 2 | 64 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Fail to cross out additional 00 at the end of test report printout. | Serial number on test report is miss read. | N/A | 1 | Operator doesn't follow procedure | 10 | Operator trained and signed off on procedure per G-QAP017. Test report is reviewed at test and pack. | 2 | 20 | N/A | N/A | N/A | - | - | - | -
|  | Default Shipping Paramters Set | Default Shipping Paramters Set | Procedure: PB560, System Final Test - P/N 10039030 | Procedures misinterpreted due lack of clarity; incorrectly written | Incorrect Shipping Parameters specified - customer recieves vent with incorrect paramters initially set | N/A | 2 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; operator trained and signed off on procedure | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow test procedure | Incorrect Shipping Parameters set - customer recieves vent with incorrect paramters initially set | N/A | 2 | Operator doesn't follow procedure - wrong vent paramters set | 10 | Operator trained and signed off on procedure per G-QAP017 | 4 | 80 | N/A | N/A | N/A | - | - | - | -
|  | Clear Vent Logs/ Faults Log Check | Clear Vent Logs/ Faults Log Check | Procedure: PB560, System Final Test - P/N 10039030 | Procedures misinterpreted due lack of clarity; incorrectly written | Faults Log is not cleared or checked properly - list of applicable faults is not complete | N/A | 2 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; Process is validated - LR0005; | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow test procedure correctly | Important fault is missed or vent is failed for normal operating fault | N/A | 2 | Operator doesn't follow procedure - fault log not cleared or important fault missed | 10 | Operator trained and signed off on procedure per G-QAP017 | 4 | 80 | N/A | N/A | N/A | - | - | - | -
|  | Pack 10039897 | Pack 10039897 | UDI Barcode Grading Verifier / equipment | Misinterpreted due to lack of clarity; incorrectly written, or operator does not follow procedure | Misinterpreted due to lack of clarity; incorrectly written, or operator does not follow procedure | N/A | 8 | Procedure not reviewed / operator error | 4 | QS00002016 formally ECO released via AGILE. Operators trained & signed off on procedure. Training matrix on line G-AMFG-1037-130 | 2 | 64 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
|  |  |  | UDI Barcode Grading Verifier / equipment | Hardware faulty, Verifier will not perform grading of Barcode | Hardware faulty, Verifier will not perform grading of Barcode | N/A | 8 | Verifier equipment malfunction | 2 | Equipment Validated per G-QAP012. Equipment Calibrated to schedule per QA G-QAP019 | 2 | 32 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
|  |  |  | Manual Operation | Label not applied, loss of tracibility | Label not applied, loss of tracibility | N/A | 7 | Operator Error | 3 | Checked at pack. Labelview allows 1 print of each serial number only | 2 | 42 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
|  | Pack 10039897 | Pack 10039897 | Process task | Incorrect s/n printed on outer carton label | FOOB | N/A | 4 | Printing error | 10 | Training; 10039897. Check in pack procedure | 1 | 40 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
|  | Pack 10039897 | Pack 10039897 | Process task | Incorrect s/n shipped | Lack of field traceability / issue at Distribution site. | N/A | 4 | Operator does not follow procedure. | 10 | Training; 10039897. Check in pack procedure | 1 | 40 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
|  |  |  | UDI Barcode Grading Verifier / equipment | Misinterpreted due to lack of clarity; incorrectly written, or operator does not follow procedure | Verifier does not grade Ship Carton Label Barcode(s) | N/A | 8 | Procedure not reviewed / operator error | 4 | QS00002016 formally ECO released via AGILE. Operators trained & signed off on procedure. Training matrix on line G-AMFG-1037-130 | 2 | 64 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
|  |  |  | UDI Barcode Grading Verifier / equipment | Hardware faulty, Verifier will not perform grading of Barcode | Operator unable to carry out step | N/A | 8 | Verifier equipment malfunction | 2 | Equipment Validated per G-QAP012. Equipment Calibrated to schedule per QA G-QAP019 | 2 | 32 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
|  |  |  | Verification Scanner | faulty/damaged | scanner does not scan barcode correctly | N/A | 8 | malfunction | 2 | PM as per G-AMFG-1180-00 | 2 | 32 | N/A | N/A | N/A
[5 empty] | Operator scans wrong label/DHR | label mismatch | N/A | 8 | Operator error | 4 | Scanner verified to detect label mismatch,PM as per G-AMFG-1180-00 | 2 | 64 | N/A | N/A | N/A
[5 empty] | Incorrect completion of DHR's | Compliance Issues. | N/A | 4 | Process instruction not followed. | 2 | Chaecked at product release. | 2 | 16 | N/A | N/A | N/A
[5 empty] | Missing labels - Dangerous good and last charge battery label | Foob; shipment held up at customs; | N/A | 5 | Process operative error | 10 | Training; 10039897. Labelling is checked per pack procedure and mainline procdure. | 1 | 50 | N/A | N/A | N/A | - | - | - | -
[5 empty] | Cosmetic issue with unit | FOOB; | N/A | 4 | Process operative error | 10 | Training; 10039897 | 2 | 80 | N/A | N/A | N/A | - | - | - | -
[5 empty] | Correct revision of labels | lack of traceability of label revision | N/A | 5 | Incomplete material clearance after change in label revision | 3 | Revision of labels controlled on Agile; revision of labels recorded on DHR for traceability; All labels will be new revision for P4.0 | 5 | 75 | N/A | N/A | N/A | - | - | - | -
[5 empty] | Process instruction | Cannot pack unit | N/A | 7 | Incomplete review | 3 | Agile review and approval | 3 | 63 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not pack Electrical Safety Test report | Customer does not receive Electrical Safety Test report | N/A | 2 | Operator: does not follow test procedure correctly | 10 | Check in DHR | 2 | 40 | N/A | N/A | N/A | - | - | - | -
|  | Software Download | Software Download | Procedure: PB560, 10039897 | Procedures misinterpreted due lack of clarity; incorrectly written | Software is not downloaded or verified to be downloaded correctly | N/A | 6 | Lack of review of procedure | 10 | Procedure is ECO controlled in Agile; process is validated - LR0005-3; Successful download message is checked when download is complete. Test process will not execute successfully without product software loaded and S/W version are recorded on DHR later in procedure | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Operator | Operator: does not follow software download procedure | Operator does not complete product software download successfully | N/A | 6 | Operator doesn't follow procedure - power switch in wrong position for ST10 download, wrong com port selected for either download app, download cable not connected | 10 | Operator trained and signed off on procedure per G-QAP017. Process is validated - LR0005-3; Successful download message is checked when download is complete. Download app cannot start with wrong com port number. Test process will not execute successfully without product software loaded | 1 | 60 | N/A | N/A | N/A | - | - | - | -
|  |  |  | Process task | Operator: does not follow process and ventilator is not assigned to correct region | Distribution centre cannot process vent as there is conflicting information on shipping carton. | N/A | 6 | Operator error. | 10 | Training; 10039897. Check in pack procedure. Check in DHR. Check on backflush log. | 2 | 120 | ALARP - Review PFMEA as part of G-QAP056. Review procedure vs practice. Check customer complaint rates. Revise occurance rating accordingly. Implementation of scanner verification check to ensure all labelling is correct. | PB560 Manufacturing Engineer. | TBD | - | - | - | -
|  | Line Clearance | Line Clearance | Operator | Parts that do not impact device functionality that are specific to PB540 or PB520 are used on PB560 | Cosmetic impact. | N/A | 4 | Incomplete line clearance | 10 | Unit is visually checked at pack and differences between PB540/560 and PB520 unique parts are easily detectable. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
|  | Line Clearance | Line Clearance | Operator | Labeling that is specific to PB540 or PB520 are used on PB560 | Incorrect indications on labeling. | N/A | 4 | Incomplete line clearance | 10 | Labeling will not impact functionality. Symbols used are common between PB540/560/520. Labeling is verified at pack. PB540/520 serial number formats will not pass PB560 functional test. | 1 | 40 | N/A | N/A | N/A | - | - | - | -
Process/Product
Failure Modes and Effects Analysis
(FMEA)

## C&E
It is not deemed necessary to complete a cause and effect matrix for the process as all the key input and output variables identified in the process flow map are considered significantly important as to be analysed by the PFMEA

## Control Plan
Control Plan
| Core Team: | [6 empty] | Date (Orig): | 2010-04-14
[8 empty] | Date (Rev): | U
Process Step | Input | Output | Process Spec (LSL, USL, Target) | Cpk / Date (Sample Size) | Measurement System | %R&R or P/T | Current Control Method (from FMEA) | Who | Where | When | Reaction Plan
N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A

## Reference Documents
Document Number | Document Title
4096600 | 560 VENTILATOR BOM
10021625 | TUBE CUTTING PROCEDURE
10021626 | LOG - PB540/PB560/PB520 TUBE CUTTING
10021754 | PB540/560/520/GEN1/POWERPACK PRODUCT RELEASE AUTHORISATION
10022945 | PB540/PB560 TEST DC POWER CABLE
10026042 | SOFTWARE REFERENCE DRAWING OPEN VENT , PB 540
10036007 | PB560 SYSTEM CALIBRATION AND TEST PROCEDURE
10038461 | PB560 DHR
10039030 | PB560, SYSTEM FINAL TEST PROCEDURE
10039032 | PB560 PEDIATRIC VOLUME AND PRESSURE TEST SCRIPT
10039033 | PB560 ADULT VOLUME AND PRESSURE TEST SCRIPT
10039034 | SOFTWARE REFERENCE DRAWING PB560 CPU MAIN
10039196 | ASSY, PB560, OUS, TOP LEVEL
10039304 | PB500 Series System Calibration/ Final Test Verification Log
10039897 | PB560 PACK PROCEDURE
10048361 | ORM, PARAMETERS LOG FOR PERIMETER SEALING MACHINE #20
10051250 | PB560 MAIN-ASSEMBLY 4 PROCEDURE.
10076002 | PB500 ELECTRICAL SAFETY TEST & BUZZER CARD CHARGE PROCEDURE
10095161 | PB560 ASSEMBLY PROCEDURE 1
10095162 | PB560 ASSEMBLY PROCEDURE 2
10095163 | PB560 ASSEMBLY PROCEDURE 3
10149038 | VERIFICATION SCANNER OPPERATING PROCEDURE
G-AMFG-1037-130 | M/S
G-AMFG-1110-00 | TEMPERATURE & HUMIDITY PROC.
G-AMFG-1180-00 | PREVENTIVE MAINTENANCE-TESTER & TEST STANDS
G-QAP004 | CHANGE CONTROL PROCEDURE
G-QAP007 | VENDOR APPROVAL
G-QAP012 | 5100 FLUKE METER
G-QAP017 | TRAINING PROCEDURE
G-QCAL019 | PERFLOW MICROMANOMETER
G-QCAL030 | TORQUE TOOL
G-QCAL041 | AIR IONISERS
G-QCAL234 | AIR IONISER GUN
II019105 | II,QUALITY STD FOR COSMETIC CRITERIA OF MECH FAB COMPONENTS
IPC-A-610 | CERTIFICATION PROGRAM FOR IPC-A-610E
QS00002016 | UDI Barcode Grading Verification procedure